In Response to Dr. Nam’s Answers on Nov. 20, 2020

Re: 	Researching and Manufacturing Standard Systems for Testing / Calibration of Dust Measuring Equipment in the Air Environment

From: 	Sibata Scientific Technology, Ltd.

To:  	Dr. Duong Thanh Nam

＜Conditions for SIBATA’s Proposals＞

A)  As Dr. Nam has previously clarified for us, the test equipment under consideration here is a system that performs "comparative tests" between standard instruments and the instruments to be tested, not a system that performs "calibration" in the sense of adjusting differences.

B)  Our detailed outlooks on this project are described below as we acknowledge your request consists of two parts: 1. to provide a standard instrument, and 2. to design and manufacture the testing system (a unit or a room?).

- In providing the standard instrument

  - The detection methods of your test target instruments are not all the same, and the accuracy of each instrument is also different. While the requested standard instrument needs to be the standard for all of these measuring devices, such an instrument are not manufactured by our company nor exists on the market.

  - We, therefore, would like to propose that you setup one standard instrument each for a real-time continuous type and a portable type. Please see the attached PowerPoint file “Proposal for MONRE_202010.” We believe there is no problem if it is operated so to be traced back to a gravimetric primary standard instrument (owned by Quatest3?).

Also (of course), we think there is no problem in using the gravimetric method for the reference primary standard indicated in the attached PowerPoint.

<Caution!>

The Instruments with low measurement concentration ranges, such as particle counters (GT-321 under No. 8 in your list, etc.), cannot be tested with the equipment we are proposing here. Separate equipment is required. It must be said that designing "a standard instrument" that can evaluate each company's products uniformly is extremely difficult, especially, for a company like ours that have little experience in manufacturing real-time meters. We will continue to discuss internally whether it is possible or not so we would appreciate it if you could give us some time.

- In designing and manufacturing the testing equipment

  - The scale of the equipment

Since you would like to divide the test target instruments into two categories, a real-time continuous type and a portable type, we recommend establishing test equipment for each. The size of the equipment will be larger for the real-time continuous type. This is because the size of the measuring instruments themselves are bigger and the drawing flow rates are higher than those of the portable types.

<Caution!>

The instruments with low measurement concentration ranges, such as particle counters (GT-321 under No. 8 in your list, etc.), cannot be tested with the equipment we are proposing here. Separate equipment is required.

  - Standard Particles

Although you indicated to use PSL, you cannot use PSL under the conditions you require. The test conditions that are not feasible for PSL are the concentration range of 0 to 1000 μg/m3 and the humidity level of 60 ± 10% RH inside the chamber. The reason is that such high concentration and high humidity make it difficult to generate PSL particles as separate single units thus making it difficult to keep particle concentration constant inside the chamber. Because hygroscopic and deliquescent particles cannot be used under the high humidity condition, it is necessary to consider particles such as silica or glass beads. If you would like a powder that imitates atmospheric dust, please let us know as we can consider it.

  - A Particle Generator

We need to design a new method of a particle generator than what we currently manufacture at our company. Considering the required concentration range, we expect our existing turntable-type particle generator will not provide sufficient concentration stability. We have several ideas of methods, but we need to make prototypes which involves trials and errors to confirm on one.

- Manufacturing time schedule and costs

As mentioned above, there are many uncertain factors that makes it difficult to determine an accurate lead time. We usually assume a period of about 5 years when conducting a product development and manufacturing of this scale. In addition, it will incur costs as we repeat making prototypes to meet your requests and make improvements to the completion. Is it possible to tell us your budget limit which you estimates?

Thank you for taking your time to consider our proposal. We hope to serve you the best we can.