Version of Record: https://www.sciencedirect.com/science/article/pii/S0269749119322730

Manuscript_e897dc8bf0edfdfb56e8ca8e687365a9

- Development of a Calibration Chamber to Evaluate the Performance of Low-cost

- Particulate Matter Sensors

- T. Sayahia, D. Kaufmanb, T. Becnelc, K. Kaura, A. Butterfielda, S. Collingwoodd, Y. Zhange,

- P.-E. Gaillardonc, K.E. Kellya

- a University of Utah, Department of Chemical Engineering, 3290 MEB, 50 S. Central Campus Dr., Salt Lake City,

- UT, United States.

- b Kaufco Consulting, LLC. 409 East Corner Bridge Lane, Draper, UT, United States.

- c University of Utah, Department of Electrical and Computer Engineering, Laboratory for NanoIntegrated Systems,

- 50 S. Central Campus Dr., Salt Lake City, UT, United States.

- d University of Utah, Department of Pediatrics, 295 Chipeta Way, Salt Lake City, UT, Unites States.

- e University of Utah, Department of Internal Medicine, Division of Epidemiology, 295 Chipeta Way, Salt Lake City,

- UT, United States.

13

# Abstract

- Low-cost particulate matter (PM) air quality sensors are becoming widely available and are

- being increasingly deployed in ambient and home/workplace environments due to their low cost,

- compactness, and ability to provide more highly resolved spatiotemporal PM concentrations.

- However, the PM data from these sensors are often of questionable quality, and the sensors need

- to be characterized individually for the environmental conditions under which they will be

- making measurements. In this study, we designed and assessed a cost-effective (~$700)

- calibration	chamber	capable	of	continuously	providing  a	uniform	PM	concentration

- simultaneously to multiple low-cost PM sensors and robust calibration relationships that are

- independent of sensor position. The chamber was designed and evaluated with a Computational

- Fluid Dynamics (CFD) model and a rigorous experimental protocol. We then used this new

- chamber to calibrate 242 Plantower PMS 3003 sensors from two production lots (Batches I and

- II) with two aerosol types: ammonium nitrate (for Batches I and II) and alumina oxide (for Batch

- I). Our CFD models and experiments demonstrated that the chamber is capable of providing

- uniform PM concentration to 8 PM sensors at once within 6% error and with excellent reliability

- (intraclass correlation coefficient > 0.771). The study identified two malfunctioning sensors and

- showed that the remaining sensors had high linear correlations with a DustTrak monitor that was

- calibrated for each aerosol type (R2>0.978). Finally, the results revealed statistically significant

- differences between the responses of Batches I and II sensors to the same aerosol (P-

- value<0.001) and the Batch I sensors to the two different aerosol types (P-value<0.001). This

- chamber design and evaluation protocol can provide a useful tool for those interested in

- systematic laboratory characterization of low-cost PM sensors.

36

# Capsule:

- Through simulations and rigorous experiments, this study designed and evaluated a cylindrical

- chamber capable of providing uniform concentrations and robust calibration relationships for

- laboratory calibration of low-cost air quality sensors.

41

42

1

© 2019 published by Elsevier. This manuscript is made available under the Elsevier user license https://www.elsevier.com/open-access/userlicense/1.0/

# Keywords:

- Air quality, Calibration chamber, Low-cost sensors, Particulate matter, Evaluation of calibration

- chamber.

46

# 1.  Introduction

- Human exposure to particulate matter (PM) is a major public health issue, contributing to lung

- cancer, cardiovascular diseases and premature mortality (Brook et al., 2010; Jensen, 2006;

- Raaschou-Nielsen et al., 2013). In 2016, ambient exposure to PM2.5 (PM with an aerodynamic

- diameter less than 2.5 μm) decreased average global life expectancy at birth by approximately

- one year with a greater reduction in life expectancy (approximately 1.2-1.9 years) in more

- polluted countries (Apte et al., 2018). PM2.5 concentration has conventionally been measured by

- expensive regulatory monitors. Although these regulatory monitors are highly accurate, their

- high costs and maintenance requirements limit the number of monitors that can be deployed.

- Consequently, these regulatory monitors are sparsely distributed and cannot capture intra-urban

- PM2.5 spatial variations resulting from local sources (Bell et al., 2010; Steinle et al., 2013).

- The emergence of low-cost light-scattering PM sensors (<$500) has provided an opportunity to

- understand neighborhood-scale differences in air pollution. A large number of such sensors have

- been deployed in various community air quality networks, such as PurpleAir (2019), AirBox

- (Chen et al., 2017) and CAIRSENSE (Jiao et al., 2016). However, data quality is a key concern

- for these PM sensors (Rai et al., 2017). They are not as accurate or precise as regulatory monitors

- (Shapiro et al., 2014). These sensors may be affected by environmental factors, such as humidity

- (Wang et al., 2015) and temperature (Gao et al., 2015). They may also experience drift over time

- (Clements et al., 2017; Sayahi et al., 2019). Many studies have relied on manufacturers’

- performance  data  rather  than  calibrating  individual  sensors  under  expected  deployment

- conditions (Ali et al., 2015; Arvind et al., 2016; Du Plessis et al., 2016; Liu et al., 2015; Van den

- Bossche et al., 2015; Zhang et al., 2017). These studies assumed that their sensors’ default

- calibration factors are appropriate for the conditions of use and that their sensors do not exhibit

- intra-sensor variability. However, several studies have demonstrated significant intra-sensor

- variability both in the field and laboratory tests	(Gao et al., 2015; Manikonda et al., 2016;

- Sayahi et al., 2019; Wang et al., 2015). Therefore, it is important to calibrate these inexpensive

- PM sensors before deployment because unreliable data may lead to unnecessary public alarm or

- complacency (Shapiro et al., 2014).

- The  performance  of  low-cost  PM  sensors  is  commonly  characterized  using  laboratory

- calibration. This type of calibration typically exposes the sensors and a research-grade/reference

- monitor to PM inside a cubic chamber. Laboratory calibration allows the sensors to experience a

- wide range of concentrations and environmental conditions in a matter of hours (Sousan et al.,

- 2016; Vercellino et al., 2018). Conversely, field calibration may require days or months, and

- even then, the sensors may not experience the target PM concentration range, PM source or

- environmental conditions. Moreover, it is logistically complicated to co-locate a large number of

- sensors with a reference monitor in the field for weeks or months.

- Calibration chambers have been a crucial tool in assessing laboratory low-cost air quality sensors

- because of their ability to provide controlled conditions (Papapostolou et al., 2017). Numerous

- studies have evaluated PM sensor performance in laboratory chambers (Table S-1). Initial

- studies took place in cubic chambers using decay tests; in these tests, an elevated particle

- concentration is reached, and then particle generation ceases. Subsequently, PM concentration

- decreases over time (Austin et al., 2015; Wang et al., 2015). Although these chambers are

- relatively inexpensive to build, they have difficulty generating stable PM concentrations, making

- it challenging to evaluate whether a low-cost PM sensor is capable of precise measurements

- because the PM concentration in the chamber is constantly dropping in a decay test. The time

- delay between the responses of PM sensors and the reference monitor is another potential source

- of error when using a decay test (Liu et al., 2017). To address the stability issue, Papapostolou et

- al. (2017) developed a sophisticated environmental chamber (cubic shaped) that can maintain a

- stable concentration over a wide range of concentrations (0-300 μg/m3 with less than 3.1%

- standard deviation).

- However, cubic chambers can have two disadvantages: particle loss at the edges and, more

- importantly, an uneven distribution of particles (or lack of understanding of this potentially

- uneven distribution). For example, Wang et al. (2015) and Hapidin et al. (2018) examined PM

- concentration uniformity by placing two low-cost sensors (same model) on different faces of

- their cubic chambers, and they suggested that PM concentration at these locations differed by

- less than 15% and 10%, respectively. However, their analysis relied on the assumption that their

- low-cost sensors’ responses exhibited no intra-sensor variability. Liu et al. (2017) developed a

- cylindrical chamber, in which the PM flows axially, and this type of chamber addresses the

- weaknesses of cubic chambers. They evaluated PM distribution in their chamber with seven

- tubes mounted at the bottom of the chamber where low-cost PM sensors are located. Sequential

- DustTrak measurements at those seven tube locations suggested that the spatial differences in

- concentration differed by less than 10%. However, this evaluation relied on the assumption that

- the PM concentration remained stable during the test. There is still a need for a cost-effective

- chamber that provides stable and uniform PM concentrations to the low-cost PM sensors as well

- as a strategy for evaluating spatial uniformity of particles within the chamber. Such a chamber

- will enable efficient evaluation and laboratory calibration of PM low-cost sensors.

- This study presents the development of a cylindrical calibration chamber that allows for the

- calibration of 8 low-cost sensors at once using continuous testing.	The distribution of PM

- concentration was examined experimentally and computationally, resulting in a robust strategy

- for assessing the spatial distribution PM concentration within the chamber. Finally, 242

- Plantower PMS 3003 PM sensors from two different purchased batches were calibrated with this

- new chamber with a goal of identifying malfunctioning sensors and understanding intra-sensor

- variability.

# 2.  Materials and methods

- This study developed and evaluated a cost-effective calibration chamber that can provide

- multiple low-cost PM sensors with a uniform PM concentration, regardless of their position in

- the chamber. The calibration system is shown in Fig. 1. This study focused primarily on the

- chamber design and used commercially available particles to evaluate the low-cost sensors.

- Several publications have discussed the aerosol generation and dilution strategies (Castell et al.,

- 2017; Li and Biswas, 2017; Liu et al., 2017; Papapostolou et al., 2017; Sousan et al., 2017). We

- used a Computational Fluid Dynamics (CFD) model to refine the chamber design before

- fabricating (described in Section 2.2).	The CFD model and a rigorous experimental and

- statistical approach were applied to evaluate the spatial variability of PM concentrations within

- the chamber.

131

132

133	Fig. 1. Schematic of the calibration system. Note the dimensions are not to scale.

134

# 2.1. Chamber design

- The new chamber, with component costs of $692 (Table S-2), is capable of evaluating eight low-

- cost PM sensors at once and has an access port for a research-grade/reference monitor. It was

- designed with several key features to encourage uniform particle concentration at the inlets of the

- low-cost sensors (Fig. 1).	These features included the cylindrical shape, the tangentially

- positioned inlet, the baffle, the outlet, and the anti-static design. The cylindrical shape limited the

- particle loss due to edge effects and provided a more uniform distribution of particles compared

- to conventional cubic chambers (Section 3.1). The inlet had turbulent flow (turbulence intensity

- of 4.51%-5.16% which falls in the medium to high turbulence case, (ANSYS, 2019) with

- incoming and exit flow velocities of 1-1.5 m/s.  The outlet of the chamber was placed in the

- center of the chamber to allow the aerosols that migrate to the edges of the chamber to flow over

- the sensors while exiting. The chamber, which was made of acrylic, also had an anti-static

- system, consisting of a helical pattern of grounded copper strips and wire (Fig. 1). An ohmmeter

- was used to verify all connections were sufficiently grounded. Generated aerosols can have high

- electrostatic charges, and the anti-static system limited particle deposits on the walls, therby

- reducing the need for cleaning (Tian et al., 2017).

151

# 2.2. Computational fluid dynamics evaluation

- CFD simulations using ANSYS FLUENT software predicted the spatial distribution of aerosols

- within the cylindrical chamber and permitted the iterative refinement of the chamber design.

- Using the same conditions and chamber volume, this study also compared CFD simulations of

- aerosol concentrations for a cubic chamber. The CFD simulations employed a realizable K-ε

- turbulence model and Eulerian-based mixture model that treats each interpenetrating phase as a

- continuous fluid, where the individual particles are not tracked. This model provides accurate

- results when the concentration of one of the phases is low and can be assumed to be well mixed.

- The primary benefit of using this model was that it had a much lower computational overhead

- than  traditional  Eulerian-Lagrangian  coupled  simulations,  where  each  particle  must  be

- individually tracked and all collisions are resolved (Tu et al., 2013). Table S-3 summarizes the

- simulation settings for the newly developed cylindrical chamber and a cubic chamber with the

- same volume.

165

# 2.3. Experimental evaluation

- The experiments included three types of tests (stability, repeatability, and rotation, described in

- Section 2.3.3) to assess the spatial distribution of PM within the chamber.  After the chamber

- evaluation, this new chamber was used to evaluate the performance of 242 low-cost sensors

- compared to a research-grade instrument (TSI DustTrak 8530) and to identify malfunctioning

- sensors as well as to understand inter-sensor variability before deploying them in the field.

172

# 2.3.1. Experimental system

- The laboratory calibration system consisted of a particle generator, a diffusion drier (if needed), a

- dilution system, and the newly designed cylindrical calibration chamber with 8-low cost PM

- sensors and a research-grade monitor (Fig. 1).

177

- Aerosol generation. Two different types of particles, alumina oxide and ammonium nitrate,

- were utilized for calibrating the low-cost sensors. Alumina oxide represented a dry and non-

- cohesive dust, and this was generated by a Palas solid particle dispenser (RBG 1000-C). The

- alumina oxide particles have a density of 4000 kg/m3 (Duralum, Washington Mills, Niagara

- Falls, NY, USA) and a mean mobility diameter of 470 nm (Fig. S-1-A). Ammonium nitrate is a

- key component of PM2.5 concentration in Salt Lake City where the sensors are being deployed.

- During winter in Salt Lake City, ammonium nitrate contributes more than 50% of PM2.5 mass

- (Baasandorj et al., 2018; Kelly et al., 2013; Sayahi et al., 2019). The 95% pure ammonium

- nitrate has a density of 1700 kg/m3 (Alfa Aesar, ThermoFisher SCIENTIFIC, ACS) and a mean

- mobility diameter of 137 nm (Fig. S-1-B). The ammonium nitrate particles were generated with a

- single-jet atomizer (TSI model 9302) and clean dry air. A building compressor provided the

- clean, dry air, which was further cleaned using an Enmet air filtration panel (5 μm, Ann Arbor,

- Michigan) that included a water collector bowl, a coalescer filter, a charcoal absorber, and an

- adsorber filter. Before introducing the aerosol flow to the calibration chamber, a diffusion dryer

- (TSI 3062) filled with dry silica gel reduced the humidity of the aerosol stream. The range of

- relative humidity inside the chamber during the calibration tests was 7.3% to 9.9%, measured

- with an Omegaette temperature/humidity meter (Omega Engineering, model HH314). The

- temperature of the chamber was similar to that of the laboratory, which ranged from 21.7 °C to

- 25.1 °C (measured with HH314). Some studies have reported that low-cost PM sensors may be

- affected by environemtal factors such as humidity (Wang et al., 2015) and temperature (Gao et

- al., 2015), but several other studies found negligible correlation between the sensors’ responses

- and those factors (Han et al., 2017; Holstius et al., 2014; Kelly et al., 2017; Sayahi et al., 2019).

- The small differences in relative humidity and temperature during the calibration tests were

- unlikely to have affected the results although as discussed in Section 3.6 the addition of

- temperature and humidity control would be beneficial for systemically evaluating these

- environmental factors.

204

- Research-grade monitor. A TSI DustTrak 8530 (Rivas et al., 2017) was placed directly below

- the calibration chamber, and a representative sample from the chamber flowed directly into the

- DustTrak (3 L/min) via conductive tubing positioned at the height of the low-cost sensors’ inlets.

- The DustTrak was equipped with a PM2.5 inlet, and its mass concentration estimate was

- corrected for each type of aerosol with a calibration factor, developed from co-located

- measurements with an AirMetrics MiniVol, which is a filter-based monitor. The MiniVol was

- operated with a PM2.5 inlet and a flow rate of 5 L/min. The MiniVol’s filter was pre- and post-

- weighed, in triplicate, to obtain an average mass concentration, which is used to develop a

- correction factor for each aerosol type. The raw DustTrak readings were divided by an

- appropriate correction factor for each aerosol type (1.78 or 0.959 for ammonium nitrate and

- alumina oxide, respectively).

216

- Low-Cost Sensors. The Plantower PM sensors (PMS 3003) were evaluated in this study (Fig. S-

- 2). In this light-scattering based sensor, a fan with a rate of 0.1 L/min routes air past a tightly

- focused laser beam. Each particle passes through two 90-degree turns before passing the laser.

- The laser wavelength is 640 ±10 nm, measured with a Lambda 35 spectrophotometer

- (PerkinElmer, Inc). The manufacturer reported that the sensor’s response time is less than 10

- seconds and its mean time to failure is more than three years. They also reported that the sensor’s

- effective PM2.5 concentration range is 0~500 μg/m3 and that the sensors have an accuracy of ±10

- μg/m3 (for concentrations between 0~100 μg/m3) and ±10% (for concentrations between 100-500

- μg/m3).

- The PMS sensor was integrated into an in-house sensing system called AirU. In addition to the

- PMS 3003 sensor, AirU comprises a custom-printed circuit board, a Texas Instruments

- HDC1080 temperature and humidity sensor, and a MiCS-4514 reduction/oxidation gas sensor.

- The PMS3003 sensors were connected directly to the AirU circuit board, which in turn was

- powered through an FCC-certified 120V AC input, 5V DC output power adapter with a 1 amp

- current rating. The PMS3003 draws 60mA, and the total current draw of the AirU (including the

- PMS3003) is 120mA. The PMS3003 sensor directly received this 5V source with minimal

- additional filtering supplied by decoupling capacitors, which were also mounted on the AirU

- circuit board. PM measurements were sent once a second from the PMS3003 to the AirU. These

- measurements were accumulated, and every 15 seconds the sum of the measurements was

- divided by the number of measurements for the time period. The result was sent to a private

- database over WiFi. No data filtering was applied prior to the data collection/storage. The PMS

- sensors were placed in a 56.1 mm diameter, circular sensor holder within the chamber. Each

- sensor was placed at 45-degree intervals along the circular holder (Fig. S-3), and the AirU boards

- were located outside the chamber (Fig. 1).

241

# 2.3.2. Calibration test protocol

- For each sensor calibration test, the calibration chamber was allowed to reach a steady-state

- target concentration based on DustTrak readings, and measurements at each concentration were

- collected for 10 min. The calibration test for all 242 low-cost sensors included 5 targets ranging

- from a PM2.5 concentration <5 μg/m3 (except the stability test, Section 2.3.3) to approximately

- 150 μg/m3 (except Section 3.5). Finally, the performance of eight sensors was also evaluated at

- 10 steady-state concentrations ranging from <5 μg/m3 to 1000 μg/m3 in a continuous test.

- Both the DustTrak and the PMS sensors collected data at 15-seconds intervals. The data were

- first time-matched and then averaged over one minute. A linear model was individually fit to the

- minute-by-minute data of DustTrak and each sensor.

252

# 2.3.3. Robustness of experimental calibration

- The following three tests provided an estimate of particle distribution inside the newly designed

- calibration chamber.

- Stability test. The stability test was performed once at 5 steady-state target concentrations (10

- min at each concentration), using DustTrak measurements (Table 1). The stability of the

- chamber at each concentration step was assessed using the standard deviation (SD) of the

- DustTrak measurements and the relative percentage of standard deviation (%RSD):

260

261

where R ef

%RSD (SD(Ref ss ))100

Ref ss

ss is the average of the DustTrak readings at each steady-state concentration.

(1)

- Repeatability test. To evaluate the ability of the chamber to generate reproducible linear

- calibration curves for each of the sensors, the 5-point calibration (described in Section 2.3.2) was

- repeated three times for eight low-cost sensors. We used the intraclass correlation coefficient

- (ICC), a widely used reliability index in test-retest analyses, to examine the repeatability

- (Cicchetti, 1994). More specifically, for each low-cost sensor, we calculated the proportion of

- residual variance after calibration, which was not caused by the repeated experiments. The

- calculated ICC ranges from 0 to 1. A high ICC index indicates a lower impact of repeated

- experiments on the calibration test and hence better repeatability in the calibration model. We

- used the cluster bootstrap approach to calculate the 95% confidence interval of ICC for each

- sensor. We then followed the guideline developed by Cicchetti (1994), and interpreted the ICC

- values less than 0.4, between 0.4 and 0.59, between 0.60 and 0.74, and greater than 0.75 as poor,

- fair, good and excellent, respectively.

- Rotation test. The rotation test complemented the repeatability test and examined the robustness

- of the calibration lines for each sensor at different positions in the chamber. In the rotation test,

- eight sensors were evaluated once at four different positions. In other words, the test was

- repeated four times, and in each test, the sensors were rotated 90-degrees in relation to the center

- of the chamber (Fig. S-3). In this test, instead of reporting the concentration differences of

- different sensors (of the same model) at various positions as an indication of the particle

- distribution (Hapidin et al., 2018; Wang et al., 2015), each sensor was compared to its own

- calibration data in the different positions. This method eliminated the errors associated with

- intra-sensor variability. We used the same ICC approach in the reliability test, discussed above,

- to examine the impact of sensor location/rotation.

284

# 2.3.4. Calibrating Plantower PMS 3003 sensors

- This study calibrated 242 Plantower PMS 3003 sensors from two different production lots

- (Batches I and II). The sensors had consecutive numbers (PMS 1- PMS 242). Batches I and II

- included 154 (PMS 1-PMS 154) and 88 sensors (PMS 155-PMS 242), respectively. They were

- purchased from Beijing Plantower Co., Ltd, in June 2017 and August 2018, respectively. Batch

- I was evaluated using both ammonium nitrate and alumina oxide aerosols; however, Batch II

- was evaluated only using ammonium nitrate particles. This calibration provided the slope,

- coefficient of determination (R2), intercept, and bias for all 242 sensors. Moreover, this study

- assessed four low-cost sensors from each batch to investigate the precision at five steady-state

- concentrations. In addition, the pairwise Spearman’s correlation coefficient (SCC) was used to

- evaluate the strength and direction of the correlation between each pair of PMS sensors for each

- test of eight sensors for both Batch I and II. The SCC ranges from -1 to 1, where the closer the

- SCC is to ±1, the stonger the association. SCC values greater than 0.80 indicate a very strong

- association (Daniel, 1990).

- It should be noted that the performance criteria for low-cost sensors do not yet exist, but other

- studies have used these same metrics (slope, intercept, R2, bias) for their evaluations of low-cost

- PM sensors (Sousan et al., 2016, 2017). In addition, the Environmental Protection Agency

- (EPA, 2006, 2018) also uses these metrics as well as the coefficient of variation (CV) in their

- criteria for continuous PM2.5 monitors and candidate equivalent methods, and we discuss these

- criteria for comparison purposes.

- The PMS sensor bias was calculated as the mean of the ratio of each PMS sensor measurement

- at time t (PMSt) to the corresponding reference reading at time t (Reft):

307

Bias  mean (PMSt	1)100

Ref

t

Reft measurements less than 1 μg/m3 were excluded from calculating sensor bias.

(2)

- The precision of the four low-cost sensors from each batch (a total of eight sensors) was assessed

- at four steady-state concentrations (PMSss) in comparison to a reference monitor using the CV:

310

311

312

313

CV (SD(PMSss ))100

Ref ss

where Ref ss is the average of reference monitor at each steady-state concentration.

# Result and discussion

  - CFD simulation

(3)

- This analysis investigated the ammonium nitrate particle concentration distribution in the

- cylindrical chamber at a flow rate of 15 L/min. The boundary conditions and the settings for the

- CFD model can be found in Table S-3. This model was tested for particle sizes up to 10 µm with

- no appreciable change in results. This CFD analysis is valid for Knudsen numbers significantly

- less than one. Fig. S-4 shows the fraction of the volume occupied by the particles in the overall

- control volume (the chamber) as time progresses. It takes the chamber about 10 minutes to reach

- steady-state. Fig. 2 represents the volume fraction profiles of the particles at the center of the

- sensors’ inlets within a ± 10 mm radial margin from the centerline (called inlet location, Fig. S-

- 5). The two smaller concentric circles in the figure represent the inlet of the DustTrack. The

- maximum concentration difference among all the sensor positions at the inlet location is 5.50%

- (Table S-4). This indicates that the cylindrical chamber is capable of providing a uniform

- concentration at the inlet location within 6% error. The difference in concentration on one side of

- the chamber is due to the inlet position and angle of the inlet. The volume fraction is lowest at

- the point where the DustTrak inlet is located. One potential reason is that the DustTrack inlet

- pulls a negative pressure while the rest of the chamber is at atmospheric pressure.

- This analysis also provided a comparison between a cubic chamber and a cylindrical one at the

- same conditions. The CFD simulation settings and boundary conditions for the cubic chamber

- are summarized in Table S-3. Fig. S-6 shows the volume fraction profiles at the six walls of the

- cubic chamber on which the sensors were assumed to be located. Comparing the concentrations

- at the center of each face of a cubic chamber showed that the smallest difference was 7.48% and

- the largest difference was 16.4% (Table S-4). Concentration differences could be even larger

- than 16.4% if one compared a center point on one face with a different location on another face

- (not the center). The results indicate that this cylindrical chamber can provide a more uniform

- distribution of particles than a cubic chamber.

338

- Fig. 2. Volume fraction distribution at the inlet location, a ±10 mm radial margin from the

- centerline of the sensors’ inlets. The two smaller concentric circles in the figure represent the

- inlet of the DustTrack. The larger concentric circle indicates the chamber outlet.

# 3.2. Effect of sensor location on calibration

- Stability, repeatability and rotation experiments were performed to evaluate the robustness of the

- sensor calibrations to position inside the newly designed chamber.

# 3.2.1.  Stability test

- Table 1, Figs. 3 and S-7 show the results of the chamber stability test for five different

- concentrations of ammonium nitrate and alumina oxide. The shaded area in the figure shows the

- transition between the steady-state concentrations. The SD and %RSD ranges for all target

- concentrations during the stability test of the new chamber were low (Table 1) and in the same

- range as those reported by Papapostolou et al. (2017) (SD=0.2-4.3, %RSD=1.4-3.1%).  These

- low SD and %RSD indicate that the chamber system has very good stability over the tested range

- of concentrations. As shown in Table 1, higher PM2.5 concentrations showed larger standard

- deviations and relative standard deviations. The same observation was reported by Papapostolou

- et al. (2017). It should be noted that %RSD is independent of concentration, making %RSD

- easier to compare with other studies (Eq. 1, Table 1).

- Papapostolou et al. (2017) discussed that the stability of PM concentrations in a chamber is

- mainly due to two factors: a) the ability of the particle generator to provide aerosol

- concentrations, and b) the uniform distribution of particles inside the chamber. In this study,

- another source of error in creating a consistent aerosol concentration could be the fluctuations in

- the clean air flow, which was provided by the building’s compressor. Our results indicated that

- the Palas RGB-C’s aerosol generation was less stable than the TSI single-jet atomizer, as

- evidenced by the larger variations in the alumina oxide tests. In order to reduce other sources of

- error and focus on the ability of the chamber to provide uniform PM concentrations, the

- repeatability and the rotation tests (Section 3.2.3) were only conducted with ammonium nitrate

- particles.

366

- Table 1. Stability test results for ammonium nitrate and alumina oxide aerosols using SD and

- %RSD (Eq. 1) parameters for five target concentrations.

|  | Ammonium nitrate Target concentration μg/m3 SD μg/m3 |  | %RSD | Alumina oxide Target concentration μg/m3 SD μg/m3 |  | %RSD |
| --- | --- | --- | --- | --- | --- | --- |
| State 1 | 8.69 | 0.506 | 5.82 | 12.5 | 0.654 | 5.57 |
| State 2 | 39.4 | 1.95 | 2.43 | 57.0 | 3.58 | 6.29 |
| State 3 | 80.2 | 2.15 | 5.45 | 96.2 | 7.73 | 8.03 |
| State 4 | 101 | 3.63 | 2.33 | 140 | 8.74 | 6.26 |
| State 5 | 156 | 4.08 | 4.06 | 175 | 9.04 | 5.17 |

- States 1-5 are the 5 steady-state concentrations for each stability test.

370

- Fig. 3. Concentration time-series for the ammonium nitrate stability tests.

# 3.2.2.  Repeatability and rotation tests

- Table 2 summarizes the ICC index values calculated for eight sensors (i.e., PMS1-PMS8, Fig. S-

- 2) compared to the DustTrak. The ICC values for repeatability and rotation tests ranged from

- 0.961 to 0.999 and from 0.771 to 0.851, respectively, which indicate excellent reliability of the

- calibration tests regardless of the repeat number or the position of the sensors inside the chamber.

- For each sensor, the ICC value for the repeatability test was consistently higher than that for the

- rotation test. The estimated uncertainty (i.e., the 95% CI) for repeatability was also consistently

- smaller than that for the rotation test. Although the impact of rotating the sensors inside the

- chamber on the linear calibration model was greater than the impact of repeating the experiment,

- the ICC index values (>0.75) of the sensors show that the calibration model results are robust to

- both the repeatability and rotation tests.

- The ICC approach was also used to describe the repeatability of each sensor’s response

- compared to itself in three repeated tests (within-sensor variability) and the repeatability of one

- sensor compared to the other seven sensors during the same test (between-sensor variability,

- three tests). The results showed that the ICC for within-sensor variability reliability ranged from

- 0.955 to 0.980, and the ICC for between-sensor variability ranged from 0.969 to 0.991. These

- ranges of ICCs are indicative of excellent reliability (Cicchetti, 1994).

389

390	Table 2. Results of repeatability and rotation tests.

# Repeatability	Rotation

391

ICC	95% CI	ICC	95% CI

| PMS1 0.987 |  | (0.938, 1.000) | 0.771 | (0.528,0.964) |
| --- | --- | --- | --- | --- |
| PMS2 0.965 |  | (0.916, 0.999) | 0.851 | (0.647,0.997) |
| PMS3 0.999 |  | (0.981, 1.000) | 0.823 | (0.555, 1.000) |
| PMS4 0.999 |  | (0.978, 1.000) | 0.809 | (0.655, 0.968) |
| PMS5 0.999 |  | (0.967, 1.000) | 0.838 | (0.649, 1.000) |
| PMS6 0.961 |  | (0.841, 1.000) | 0.833 | (0.666, 0.998) |
| PMS7 0.985 |  | (0.928, 1.000) | 0.790 | (0.597, 1.000) |
|  | PMS8 0.995 (0.978, 1.000) 0.843 (0.595, 0.977) |  |  |  |
| 392 |  |  |  |  |
| 393 | 3.3. Effect of aerosol type |  |  |  |
| 394 | The evaluation of the PMS 3003 performance for two different particle types (alumina oxide and |  |  |  |

- ammonium nitrate) began with a PM2.5 time-series evaluation of the DustTrak and four PMS

- sensors at five steady-state concentrations (Fig. S-8). The figure illustrates that the PMS sensors

- track the DustTrak PM2.5 concentrations. As shown in Fig. S-8-A, the PMS sensors from Batch I

- overestimated  ammonium  nitrate  PM2.5  concentrations  compared  to  the  DustTrak  but

- underestimated alumina oxide concentrations compared to DustTrak (Fig. S-8-B). Kelly et al.

- (2017) also reported that the PMS 3003 underestimated alumina oxide PM2.5 concentrations in a

- wind-tunnel test. Many studies suggested that the differences in the response of light-scattering

- based sensors to different aerosol types can be attributed to variations in aerosol optical

- properties (Kelly et al., 2017; Liu et al., 2017; Sousan et al., 2016; Wang et al., 2015).

- A linear model was fit to each of the 242 sensors based on the 5-point calibration curve. Fig. 4

- shows the distribution of the slopes, intercepts, and goodness of fit metrics (R2) of these fits as

- well as bias measurements. For both aerosol types, the PM2.5 readings of Batch I PMS sensors

- were highly correlated with the reference instrument, R2 of 0.995-0.999 for ammonium nitrate

- and R2 of 0.978-0.998 for alumina oxide. Clements et al. (2017) suggested that low-cost sensors

- can complement existing monitoring networks and enhance the spatial coverage of PM data if

- they have R2>0.4 with a federal reference/equivalent method. As discussed in Section 2.3.4,

- performance standards for low-cost sensors do not yet exist; for comparison purposes, we

- discussed EPA criteria for continuous PM2.5 monitors and candidate equivalent methods. The

- correlations between the PMS sensors and the DustTrak meet the EPA’s criteria for continuous

- PM2.5 monitors (r>0.9 or R2>0.81; EPA, 2018) and for candidate equivalent methods (r>0.97 or

- R2>0.94; EPA, 2006). However, the slopes of the linear model were not generally within the

- EPA’s criteria for continuous PM2.5 monitors (1 ± 0.1; EPA, 2018) and varied significantly for

- the  two  different  types  of  aerosols  (Student  t-test,  p-value<0.001)  from  0.414-0.645

- (underestimation for alumina oxide) to 0.988-1.36 (overestimation for ammonium nitrate). The

- average of intercepts of all Batch I sensors were -1.37 μg/m3 (range of -3.88-2.47 μg/m3) and -

- 1.94 μg/m3 (range of -5.84-5.05 μg/m3) for ammonium nitrate and alumina oxide, respectively.

- The small intercepts for the Batch I sensors for both aerosol types were generally within the EPA

- (2006) acceptable range for candidate equivalent methods (± 5 μg/m3).

- The average bias values for all the Batch I sensors for ammonium nitrate and alumina oxide were

- 9.53% and -49.6%, respectively. The small positive bias for Batch I sensors with ammonium

- nitrate shows that the difference between the sensor and the reference instrument is small and

- within the EPA (2018) criteria for continuous PM2.5 monitors (± 10%). However, the large

- negative bias for alumina oxide exceeds EPA (2018) criteria.  It is also an indication of the

- underestimation of PM2.5 concentrations. Sousan et al. (2017) also reported different bias values

- for the same sensor (Footbot) in response to different aerosol types (-12% for Arizona air dust

- and < -46% for salt and welding fume).

431

(A)

(B)

- Fig. 4. Frequency distribution of slopes, intercepts, coefficients of determination, and bias for

- 154 Batch I sensors for calibration with ammonium nitrate (A) and alumina oxide (B) aerosols.

- Note that the scales in the x-axes are different.

- The precisions of the four representative PMS sensors were between 1.96% and 10.2% (average

- of 5.06%) for ammonium nitrate and between 1.17% to 5.91% (average of 3.65%) for alumina

- oxide. The low precision values reflect the ability of the sensors in replicating the concentration

- measurements (NIOSH, 2012). These CVs are within EPA’s acceptable measurement

- uncertainty for continuous PM2.5 monitors (CV < 10%, Table S-5; EPA, 2018).

# 3.4. Comparison of Batches I and II sensors

- Comparing the response of Batches I and II sensors revealed that Batch I slightly overestimated

- ammonium nitrate PM2.5 concentrations (mean slope = 1.16 and bias = 9.53 %), but Batch II

- slightly underestimated ammonium nitrate PM2.5 concentrations (mean slope = 0.919 and bias = -

- 4.52 %, Fig. S-9). As with Batch I, the Batch II sensor responses were highly correlated with the

- DustTrak (R2 of 0.968-0.999). These slopes and correlations meet EPA (2018) criteria for

- continuous PM2.5 monitors. All of these results in the following sections excluded the two

- malfunctioning sensors, which are further discussed below. As shown in Fig. S-10, the slopes of

- Batch I sensors differed significantly from those of Batch I (Student t-test, p-value<0.001). This

- result indicates that sensors purchased at different times could have different responses.

- However, the pairwise SCCs showed very strong positive correlation for each pair of eight

- sensors in the same test (>0.972 for Batch I and >0.978 for Batch II). Table S-6 shows the

- pairwise SCCs for one representative test from each batch.

- Eight sensors from Batch II have much lower slopes than the mean slopes (0.5 to 0.7) and biases

- (-50 % to -20 %) although they still have high R2 values (0.995-0.997). This means that these

- sensors underestimate the PM2.5 levels more than the other sensors. However, they could still be

- deployed in the sensor network using proper correction factors. These results provide indications

- of intra-sensor variability and the importance of calibration before field deployment of the

- sensors. Possible explanations for the different responses of the batches include manufacturer

- modification of their software that converts the signal to PM concentration or manufacturing

- differences in the photodetector.

- One of the goals of laboratory evaluation was to identify defective sensors and avoid deploying

- them in the field. The malfunctioning sensors were identified by a poor coefficient of

- determination (R2 < 0.3, Moore et al., 2013). Two out of 242 evaluated sensors were classified as

- malfunctioning (PMS 163 and PMS 242). Neither sensor tracked the PM2.5  concentration

- measurements of DustTrak. Sensor PMS 163 readings ranged from 0-24.3 μg/m3 with R2 of

- 0.018 (Fig. S-11) while sensor PMS 242 constantly read zero with R2 of NaN (not a number),

- even though both sensors were exposed to PM2.5 concentrations up to 150 μg/m3. Both sensors

- were dismantled to find possible reasons for their behavior. Nothing suspicious was found in

- sensor PMS 163; its low readings could be due to laser misalignment or a manufacturing defect

- in the sensor’s laser, photodetector or electronics. Sensor PMS 242’s laser chamber was filled

- with small pieces of white paper (Fig. S-12) preventing the laser light from reaching the particles

- or the photodetector.

473

# 3.5. Sensor behavior at higher concentrations

- Exposing eight PMS sensors to ammonium nitrate PM2.5 concentrations up to 1000 μg/m3

- showed that the sensors maintained a linear response up to a concentration of 150 μg/m3 (R2 of

- 0.995-0.999, Fig. S-13). However, the linearity deteriorated at higher concentrations. The exact

- physical reason behind the non-linear relationship of PMS 3003 sensors and the reference

- instrument is unclear. A second-order polynomial fit all the data (R2 of 0.998-0.999, range: 5 -

- 1000 μg/m3) better than a linear model (Fig S-14). This non-linear behavior of low-cost PM

- sensors above a certain concentration was also seen in other laboratory studies (Austin et al.,

- 2015; Hapidin et al., 2018; Sousan et al., 2017; Wang et al., 2015).

# 3.6. Limitations

- The focus of this paper was to develop, evaluate, and demonstrate a cost-effective cylindrical

- chamber for calibration of low-cost PM sensors. A number of refinements would enhance the

- usability of this system. For example, using this chamber to calibrate hundreds of PM sensors

- can be time-consuming. An improved experimental design could streamline the calibration

- procedure and/or potentially reduce the number of data points needed for sensor calibration.

- Adding a neutralizer after the aerosol generation system could prevent particle deposits in the

- lines prior to the chamber; however, this would significantly add to the cost. Using copper mesh

- instead of copper strip in the anti-static system is another potential improvement. It would also

- be beneficial to integrate temperature and humidity controls into the system for future studies

- because some investigations have shown that temperature and relative humidity can affect the

- responses of some low-cost sensors (Gao et al., 2015; Wang et al., 2015).

- A laboratory evaluation system is well suited for calibration of sensors prior to field deployment,

- and it can complement, but not replace, field calibration. Aerosols in the real word are

- heterogenous, and their properties vary over time. A one-time lab calibration cannot represent

- the range of real-world conditions that a sensor may experience, and it cannot ensure that

- measurements are accurate in the real world. However, it does allow one to evaluate low-cost

- sensor performance (i.e., intra-sensor variability, response over relevant concentration ranges),

- and aids with quality assurance and quality contorl concerns, such as identifying malfunctioning

- sensors. A mixture of lab and field calibrations along with advanced statistical techniques is a

- likely path forward (Johnson et al., 2018; Becnel et al., 2019.; Zimmerman et al., 2018).

# 4.  Conclusion

- This study developed and evaluated a cost-effective calibration chamber that provided spatially

- uniform concentrations to multiple PM sensors and calibration relationships that were robust to

- sensor position in the chamber. The design and evaluation relied on a CFD model and a rigorous

- experimental evaluation. The computational model showed that the chamber is capable of

- providing a uniform PM concentration to calibrate eight sensors at one time within 6% error, and

- the experimental results demonstrated the robustness of the calibration model to the sensor

- position within the chamber with excellent reliability (ICC>0.771). This new chamber was then

- used to evaluate the performance of 242 Plantower PMS 3003 sensors from two production lots

- using two different particle types: ammonium nitrate (for Batches I and II) and alumina oxide

- (for Batch I). The results identified two malfunctioning sensors and demonstrated that all the

- sensors (except the two malfunctioning sensors) were highly correlated with the DustTrak

- reference monitor (R2>0.978). The study also identified significant response differences between

- Batches I and II. As in several other studies, the PMS sensors exhibited a statistically significant

- difference in their responses to the two different aerosol types. This calibration chamber was able

- to identify manufacturer differences in PM response and malfunctioning sensors. This chamber

- can also be used to complement field observations and help to eventually determine whether

- good laboratory performance of a low-cost PM sensor indicates its good performance under real-

- world conditions.

523

# Acknowledgment

- We  gratefully  acknowledge  National  Science  Foundation  (NSF)  support  under  award

- numbers 1646408 and 1642513. Research reported in this paper was also supported in part by the

- ECHO Program, National Institutes of Health under Award Number UH3OD023249, the

- PRISMS Program, National Institute of Biomedical Imaging and Bioengineering of the National

- Institutes of Health under Award Number U54EB021973 and the University of Utah Program for

- Air Quality, Health and Society. The content is solely the responsibility of the authors and does

- not necessarily represent the official views of the sponsors.

532

# References

- Ali, H., Soe, J.K., Weller, S.R., 2015. A real-time ambient air quality monitoring wireless sensor

- network for schools in smart cities. 2015 IEEE 1st Int. Smart Cities Conf. ISC2 2015 5–10.

- doi:10.1109/ISC2.2015.7366163

- ANSYS, 2019. ANSYS FLUENT 12.0 User’s Guide - 7.3.2 Using Flow Boundary Conditions

- [WWW Document]. URL

- http://www.afs.enea.it/project/neptunius/docs/fluent/html/ug/node238.htm (accessed

- 4.30.19).

- Apte, J.S., Brauer, M., Cohen, A.J., Ezzati, M., Pope, C.A., 2018. Ambient PM2.5Reduces

- Global and Regional Life Expectancy. Environ. Sci. Technol. Lett. 5, 546–551.

- doi:10.1021/acs.estlett.8b00360

- Arvind, D.K., Mann, J., Bates, A., Kotsev, K., 2016. The AirSpeck Family of Static and Mobile

- Wireless Air Quality Monitors. Proc. - 19th Euromicro Conf. Digit. Syst. Des. DSD 2016

- 207–214. doi:10.1109/DSD.2016.110

- Austin, E., Novosselov, I., Seto, E., Yost, M.G., 2015. Laboratory evaluation of the Shinyei

- PPD42NS low-cost particulate matter sensor. PLoS One 10, 1–17.

- doi:10.1371/journal.pone.0137789

- Baasandorj, M., Brown, S., Hoch, S., Crosman, E., Long, R., Silva, P., Mitchell, L., Hammond,

- I., Martin, R., Bares, R., Lin, J., Sohl, J., Page, J., McKeen, S., Pennell, C., Franchin, A.,

- Petersen, R., Hallar, G., Fibiger, D., Womack, C., McDuffie, E., Moravek, A., Murphy, J.,

- Hrdina, A., Thornton, J., Goldberger, L., Lee, B., Riedel, T., Whitehill, A., Kelly, K.,

- Hansen, J., Eatough, D., 2018. 2017 Utah Winter Fine Particulate Study Final Report.

- Becnel, T., Sayahi, T., Kelly, K., Gaillardon, P., 2019. A Recursive Approach to Partially Blind

- Calibration of a Pollution Sensor Network. The 15th IEEE International Conference on

- Embedded Software and Systems, Las Vegas.

- Bell, M.L., Ebisu, K., Peng, R.D., 2010. Community-level spatial heterogeneity of chemical

- constituent levels of fine particulates and implications for epidemiological research. J.

- Expo. Sci. Environ. Epidemiol. 21, 372–384. doi:10.1038/jes.2010.24

- Brook, R.D., Rajagopalan, S., Iii, C.A.P., Brook, J.R., Bhatnagar, A., Diez-roux, A. V, Holguin,

- F., Hong, Y., Luepker, R. V, Mittleman, M.A., Peters, A., Siscovick, D., Smith, S.C.,

- Whitsel, L., Kaufman, J.D., 2010. Particulate matter air pollution and cardiovascular disease

- an update to the scientific statement from the american.

- doi:10.1161/CIR.0b013e3181dbece1

- Castell, N., Dauge, F.R., Schneider, P., Vogt, M., Lerner, U., Fishbain, B., Broday, D.,

- Bartonova, A., 2017. Can commercial low-cost sensor platforms contribute to air quality

- monitoring and exposure estimates? Environ. Int. 99, 293–302.

- doi:10.1016/j.envint.2016.12.007

- Chen, L.-J., Ho, Y.-H., Lee, H.-C., Wu, H.-C., Liu, H.-M., Hsieh, H.-H., Huang, Y.-T., Lung, S.-

- C.C., 2017. An Open Framework for Participatory PM2.5 Monitoring in Smart Cities. IEEE

- Access 5, 14441–14454. doi:10.1109/ACCESS.2017.2723919

- Cicchetti, D. V, 1994. Guidlines, Criteria, and Rules of Thumb for Evalauting Normed and

- Standardized Assessment Instruments in Psychology. Psychol. Assess. 6, 284–290.

- doi:10.1037/1040-3590.6.4.284

- Clements, A.L., Griswold, W.G., RS, A., Johnston, J.E., Herting, M.M., Thorson, J., Collier-

- Oxandale, A., Hannigan, M., 2017. Low-Cost Air Quality Monitoring Tools: From

- Research to Practice (A Workshop Summary). Sensors 17, 2478. doi:10.3390/s17112478

- Daniel, W.W., 1990. Applied nonparametric statistics, 2nd ed. PWS-KENT Pub.

- Du Plessis, R., Kumar, A., Hancke, G.P., Silva, B.J., 2016. A wireless system for indoor air

- quality monitoring. IECON Proc. (Industrial Electron. Conf. 5409–5414.

- doi:10.1109/IECON.2016.7794087

- EPA, 2018. Technical Note - PM 2.5 Continuous Monitor Comparability Assessment 1.

- EPA, 2006. 40 CFR Appendix Table_C-4_to_subpart_C_of_part_53 - Test Specifications for

- PM 10, PM 2.5 and PM 10-2.5 Candidate Equivalent Methods | CFR | US Law | LII / Legal

- Information Institute [WWW Document]. URL

- https://www.law.cornell.edu/cfr/text/40/appendix-Table_C-4_to_subpart_C_of_part_53

- (accessed 4.10.19).

- Gao, M., Cao, J., Seto, E., 2015. A distributed network of low-cost continuous reading sensors to

- measure spatiotemporal variations of PM2.5 in Xi’an, China. Environ. Pollut. 199, 56–65.

- doi:10.1016/j.envpol.2015.01.013

- Han, I., Symanski, E., Stock, T.H., 2017. Feasibility of using low-cost portable particle monitors

- for measurement of fine and coarse particulate matter in urban ambient air. J. Air Waste

- Manage. Assoc. 67, 330–340. doi:10.1080/10962247.2016.1241195

- Hapidin, D.A., Saputra, C., Maulana, D.S., Munir, M.M., Khairurrijal, K., 2018. Aerosol

- Chamber Characterization for Commercial Particulate Matter (PM) Sensor Evaluation.

- Aerosol Air Qual. Res. 181–194. doi:10.4209/aaqr.2017.12.0611

- Holstius, D.M., Pillarisetti, A., Smith, K.R., Seto, E., 2014. Field calibrations of a low-cost

- aerosol sensor at a regulatory monitoring site in California. Atmos. Meas. Tech. 7, 1121–

- 1131. doi:10.5194/amt-7-1121-2014

- Jensen, M.D., 2006. Adipose tissue as an endocrine organ: Implications of its distribution on free

- fatty acid metabolism. Eur. Hear. Journal, Suppl. 8. doi:10.1038/nature15371

- Jiao, W., Hagler, G., Williams, R., Sharpe, R., Brown, R., Garver, D., Judge, R., Caudill, M.,

- Rickard, J., Davis, M., Weinstock, L., Zimmer-dauphinee, S., Buckley, K., 2016.

- Community Air Sensor Network ( CAIRSENSE ) project : evaluation of low-cost sensor

- performance in a suburban environment in the southeastern United States 5281–5292.

- doi:10.5194/amt-9-5281-2016

- Johnson, N.E., Bonczak, B., Kontokosta, C.E., 2018. Using a gradient boosting model to

- improve the performance of low-cost aerosol monitors in a dense, heterogeneous urban

- environment. Atmos. Environ. 184, 9–16. doi:10.1016/j.atmosenv.2018.04.019

- Kelly, K.E., Kotchenruther, R., Kuprov, R., Silcox, G.D., 2013. Receptor model source

- attributions for Utah’s Salt Lake City airshed and the impacts of wintertime secondary

- ammonium nitrate and ammonium chloride aerosol. J. Air Waste Manag. Assoc. 63, 575–

- 590. doi:10.1080/10962247.2013.774819

- Kelly, K.E., Whitaker, J., Petty, A., Widmer, C., Dybwad, A., Sleeth, D., Martin, R., Butterfield,

- A., 2017. Ambient and laboratory evaluation of a low-cost particulate matter sensor.

- Environ. Pollut. 221, 491–500. doi:10.1016/j.envpol.2016.12.039

- Li, J., Biswas, P., 2017. Optical Characterization Studies of a Low-Cost Particle Sensor. Aerosol

619	Air Qual. Res. 17, 1691–1704. doi:10.4209/aaqr.2017.02.0085

- Liu, D., Zhang, Q., Jiang, J., Chen, D.R., 2017. Performance calibration of low-cost and portable

- particular matter (PM) sensors. J. Aerosol Sci. 112, 1–10.

- doi:10.1016/j.jaerosci.2017.05.011

- Liu, X., Li, B., Jiang, A., Qi, S., Xiang, C., Xu, N., 2015. A bicycle-borne sensor for monitoring

- air pollution near roadways. 2015 IEEE Int. Conf. Consum. Electron. - Taiwan, ICCE-TW

- 2015 166–167. doi:10.1109/ICCE-TW.2015.7216835

- Manikonda, A., Zikova, N., Hopke, P.K., Ferro, A.R., 2016. Laboratory assessment of low-cost

627	PM monitors. J. Aerosol Sci. 102, 29–40. doi:10.1016/j.jaerosci.2016.08.010

- Moore, D., Notz, W., Fligner, M., 2013. The basic practice of statistics, 6th ed. W. H. Freeman

- and Company, New York.

- National Institute of Occupational Safety and Health, 2012. Components for Evaluation of

- Direct-Reading Monitors for Gases and Vapors.

- Papapostolou, V., Zhang, H., Feenstra, B.J., Polidori, A., 2017. Development of an

- environmental chamber for evaluating the performance of low-cost air quality sensors under

- controlled conditions. Atmos. Environ. 171, 82–90. doi:10.1016/j.atmosenv.2017.10.003

- PurpleAir, 2019. PurpleAir Map, Air Quality Map [WWW Document]. URL

- http://map.purpleair.org/. (accessed 3.19.19).

- Raaschou-Nielsen, O., Andersen, Z.J., Beelen, R., Samoli, E., Stafoggia, M., Weinmayr, G.,

- Hoffmann, B., Fischer, P., Nieuwenhuijsen, M.J., Brunekreef, B., Xun, W.W., Katsouyanni,

- K., Dimakopoulou, K., Sommar, J., Forsberg, B., Modig, L., Oudin, A., Oftedal, B.,

- Schwarze, P.E., Nafstad, P., De Faire, U., Pedersen, N.L., Östenson, C.-G., Fratiglioni, L.,

- Penell, J., Korek, M., Pershagen, G., Eriksen, K.T., Sørensen, M., Tjønneland, A.,

- Ellermann, T., Eeftens, M., Peeters, P.H., Meliefste, K., Wang, M., Bueno-de-Mesquita, B.,

- Key, T.J., de Hoogh, K., Concin, H., Nagel, G., Vilier, A., Grioni, S., Krogh, V., Tsai, M.-

- Y., Ricceri, F., Sacerdote, C., Galassi, C., Migliore, E., Ranzi, A., Cesaroni, G., Badaloni,

- C., Forastiere, F., Tamayo, I., Amiano, P., Dorronsoro, M., Trichopoulou, A., Bamia, C.,

- Vineis, P., Hoek, G., 2013. Air pollution and lung cancer incidence in 17 European cohorts:

- prospective analyses from the European Study of Cohorts for Air Pollution Effects

- (ESCAPE). Lancet Oncol. 14, 813–822. doi:10.1016/S1470-2045(13)70279-1

- Rai, A.C., Kumar, P., Pilla, F., Skouloudis, A.N., Di Sabatino, S., Ratti, C., Yasar, A., Rickerby,

- D., 2017. End-user perspective of low-cost sensors for outdoor air pollution monitoring.

- Sci. Total Environ. 607–608, 691–705. doi:10.1016/j.scitotenv.2017.06.266

- Rivas, I., Mazaheri, M., Viana, M., Moreno, T., Clifford, S., He, C., Bischof, O.F., Martins, V.,

- Reche, C., Alastuey, A., Alvarez-Pedrerol, M., Sunyer, J., Morawska, L., Querol, X., 2017.

- Identification of technical problems affecting performance of DustTrak DRX aerosol

- monitors. Sci. Total Environ. 584–585, 849–855. doi:10.1016/J.SCITOTENV.2017.01.129

- Sayahi, T., Butterfield, A., Kelly, K.E., 2019. Long-term field evaluation of the Plantower PMS

- low-cost particulate matter sensors *. doi:10.1016/j.envpol.2018.11.065

- Shapiro, M.A., Swann, P.G., Hartsough, M., 2014. Regulatory Considerations of Lower Cost Air

- Pollution Sensor Data Performance. Handb. Ther. Antibodies 1, 277–300.

- doi:10.1002/9783527619740.ch12

- Sousan, S., Koehler, K., Hallett, L., Peters, T.M., 2017. Evaluation of consumer monitors to

- measure particulate matter. J. Aerosol Sci. 107, 123–133.

- doi:10.1016/j.jaerosci.2017.02.013

- Sousan, S., Koehler, K., Thomas, G., Park, J.H., Halterman, A., Peters, T.M., Sousan, S.,

- Koehler, K., Thomas, G., Park, J.H., Hillman, M., 2016. Inter-comparison of low-cost

- sensors for measuring the mass concentration of occupational aerosols. Aerosol Sci.

- Technol. 50, 462–473. doi:10.1080/02786826.2016.1162901

- Steinle, S., Reis, S., Sabel, C.E., 2013. Quantifying human exposure to air pollution—Moving

- from static monitoring to spatio-temporally resolved personal exposure assessment. Sci.

- Total Environ. 443, 184–193. doi:10.1016/J.SCITOTENV.2012.10.098

- Tian, J., Brem, B.T., West, M., Bond, T.C., Rood, M.J., Riemer, N., 2017. Simulating aerosol

- chamber experiments with the particle-resolved aerosol model PartMC. Aerosol Sci.

- Technol. 51, 856–867. doi:10.1080/02786826.2017.1311988

- Tu, J., Yeoh, G.H., Liu, C., 2013. Computational fluid dynamics : a practical approach, second.

- ed. Butterworth-Heinemann.

- Van den Bossche, J., Peters, J., Verwaeren, J., Botteldooren, D., Theunis, J., De Baets, B., 2015.

- Mobile monitoring for mapping spatial variation in urban air quality: Development and

- validation of a methodology based on an extensive dataset. Atmos. Environ. 105, 148–161.

- doi:10.1016/j.atmosenv.2015.01.017

- Vercellino, R.J., Sleeth, D.K., Handy, R.G., Min, K.T., Collingwood, S.C., 2018. Laboratory

- evaluation of a low-cost, real-time, aerosol multi-sensor. J. Occup. Environ. Hyg. 15, 559–

- 567. doi:10.1080/15459624.2018.1468565

- Wang, Y., Li, J., Jing, H., Zhang, Q., Jiang, J., Biswas, P., 2015. Laboratory Evaluation and

- Calibration of Three Low-Cost Particle Sensors for Particulate Matter Measurement.

- Aerosol Sci. Technol. 49, 1063–1077. doi:10.1080/02786826.2015.1100710

- Zhang, R., Ravi, D., Yang, G.-Z., Lo, B., 2017. A personalized air quality sensing system - a

- preliminary study on assessing the air quality of London underground stations. 2017 IEEE

- 14th Int. Conf. Wearable Implant. Body Sens. Networks 111–114.

- doi:10.1109/BSN.2017.7936020

- Zimmerman, N., Presto, A.A., Kumar, S.P.N., Gu, J., Hauryliuk, A., Robinson, E.S., Robinson,

- A.L., 2018. A machine learning calibration model using random forests to improve sensor

- performance for lower-cost air quality monitoring. Atmos. Meas. Tech. 11, 291–313.

- doi:10.5194/amt-11-291-2018

694

695

Graphical abstract