## 001
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 001 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103770 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.3 |  |  |  |  | ℃ |  | 99.3 |  |  | ℃ |  | 0 |  |  | ℃
| 200℃ |  |  | 199.6 |  |  |  |  | ℃ |  | 198 |  |  | ℃ |  | -1.6 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 002
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 002 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103771 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 98.8 |  |  |  |  | ℃ |  | 98.8 |  |  | ℃ |  | 0 |  |  | ℃
| 200℃ |  |  | 198.5 |  |  |  |  | ℃ |  | 198.8 |  |  | ℃ |  | 0.3 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 003
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 003 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103772 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 98.8 |  |  |  |  | ℃ |  | 98.1 |  |  | ℃ |  | -0.7 |  |  | ℃
| 200℃ |  |  | 198.5 |  |  |  |  | ℃ |  | 198 |  |  | ℃ |  | -0.5 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 004
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 004 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103773 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 98.8 |  |  |  |  | ℃ |  | 99.2 |  |  | ℃ |  | 0.4 |  |  | ℃
| 200℃ |  |  | 198.5 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 1.5 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 005
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 005 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103774 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.4 |  |  |  |  | ℃ |  | 99.4 |  |  | ℃ |  | 0 |  |  | ℃
| 200℃ |  |  | 202.6 |  |  |  |  | ℃ |  | 203 |  |  | ℃ |  | 0.4 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 006
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 006 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103775 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 98.8 |  |  |  |  | ℃ |  | 98.8 |  |  | ℃ |  | 0 |  |  | ℃
| 200℃ |  |  | 198.5 |  |  |  |  | ℃ |  | 199 |  |  | ℃ |  | 0.5 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 007
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 007 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103776 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 98.5 |  |  |  |  | ℃ |  | 99.6 |  |  | ℃ |  | 1.1 |  |  | ℃
| 200℃ |  |  | 199 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 1 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 008
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 008 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103777 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 98.5 |  |  |  |  | ℃ |  | 97.5 |  |  | ℃ |  | -1 |  |  | ℃
| 200℃ |  |  | 199 |  |  |  |  | ℃ |  | 199 |  |  | ℃ |  | 0 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 009
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 009 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103778 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.2 |  |  |  |  | ℃ |  | 98.5 |  |  | ℃ |  | -0.7 |  |  | ℃
| 200℃ |  |  | 199 |  |  |  |  | ℃ |  | 199 |  |  | ℃ |  | 0 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 010
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 010 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103779 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.2 |  |  |  |  | ℃ |  | 98 |  |  | ℃ |  | -1.2 |  |  | ℃
| 200℃ |  |  | 199 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 1 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 011
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 011 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103780 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99 |  |  |  |  | ℃ |  | 97.2 |  |  | ℃ |  | -1.8 |  |  | ℃
| 200℃ |  |  | 199 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 1 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 012
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 012 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103781 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 100.9 |  |  |  |  | ℃ |  | 100.5 |  |  | ℃ |  | -0.4 |  |  | ℃
| 200℃ |  |  | 199.3 |  |  |  |  | ℃ |  | 199 |  |  | ℃ |  | -0.3 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 013
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 013 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103752 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 98.5 |  |  |  |  | ℃ |  | 98.7 |  |  | ℃ |  | 0.2 |  |  | ℃
| 200℃ |  |  | 199 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 1 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 014
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 014 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103753 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 100.6 |  |  |  |  | ℃ |  | 100.4 |  |  | ℃ |  | -0.2 |  |  | ℃
| 200℃ |  |  | 199.3 |  |  |  |  | ℃ |  | 199 |  |  | ℃ |  | -0.3 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 015
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 015 | [8 empty] | Calibration day |  | 2017-11-01
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103754 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 101 |  |  |  |  | ℃ |  | 100.4 |  |  | ℃ |  | -0.6 |  |  | ℃
| 200℃ |  |  | 199.3 |  |  |  |  | ℃ |  | 198 |  |  | ℃ |  | -1.3 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 016
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 016 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103755 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 100.6 |  |  |  |  | ℃ |  | 100.6 |  |  | ℃ |  | 0 |  |  | ℃
| 200℃ |  |  | 199.3 |  |  |  |  | ℃ |  | 199 |  |  | ℃ |  | -0.3 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 017
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 017 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103756 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 100.8 |  |  |  |  | ℃ |  | 100 |  |  | ℃ |  | -0.8 |  |  | ℃
| 200℃ |  |  | 199.3 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 0.7 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 018
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 018 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103757 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 100.8 |  |  |  |  | ℃ |  | 99.9 |  |  | ℃ |  | -0.9 |  |  | ℃
| 200℃ |  |  | 199.3 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 0.7 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 019
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 019 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103758 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99 |  |  |  |  | ℃ |  | 98.6 |  |  | ℃ |  | -0.4 |  |  | ℃
| 200℃ |  |  | 199.4 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 0.6 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 020
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 020 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103759 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.1 |  |  |  |  | ℃ |  | 99.7 |  |  | ℃ |  | 0.6 |  |  | ℃
| 200℃ |  |  | 199.2 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 0.8 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 021
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 021 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103760 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.1 |  |  |  |  | ℃ |  | 99.8 |  |  | ℃ |  | 0.7 |  |  | ℃
| 200℃ |  |  | 199.8 |  |  |  |  | ℃ |  | 201 |  |  | ℃ |  | 1.2 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 022
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 022 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103761 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 98.9 |  |  |  |  | ℃ |  | 98.9 |  |  | ℃ |  | 0 |  |  | ℃
| 200℃ |  |  | 199.4 |  |  |  |  | ℃ |  | 201 |  |  | ℃ |  | 1.6 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 023
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 023 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103762 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.2 |  |  |  |  | ℃ |  | 99.2 |  |  | ℃ |  | 0 |  |  | ℃
| 200℃ |  |  | 199.8 |  |  |  |  | ℃ |  | 201 |  |  | ℃ |  | 1.2 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 024
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 024 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103763 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 101.8 |  |  |  |  | ℃ |  | 100.7 |  |  | ℃ |  | -1.1 |  |  | ℃
| 200℃ |  |  | 199.3 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 0.7 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 025
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 025 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103764 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 98.9 |  |  |  |  | ℃ |  | 98.9 |  |  | ℃ |  | 0 |  |  | ℃
| 200℃ |  |  | 198.5 |  |  |  |  | ℃ |  | 199 |  |  | ℃ |  | 0.5 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 026
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 026 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103765 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.1 |  |  |  |  | ℃ |  | 98.8 |  |  | ℃ |  | -0.3 |  |  | ℃
| 200℃ |  |  | 199.8 |  |  |  |  | ℃ |  | 201 |  |  | ℃ |  | 1.2 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 027
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 027 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103766 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.1 |  |  |  |  | ℃ |  | 99.1 |  |  | ℃ |  | 0 |  |  | ℃
| 200℃ |  |  | 199.8 |  |  |  |  | ℃ |  | 201 |  |  | ℃ |  | 1.2 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 028
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 028 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103767 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.2 |  |  |  |  | ℃ |  | 98.9 |  |  | ℃ |  | -0.3 |  |  | ℃
| 200℃ |  |  | 199.2 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 0.8 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 029
o | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 029 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103768 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.2 |  |  |  |  | ℃ |  | 98.9 |  |  | ℃ |  | -0.3 |  |  | ℃
| 200℃ |  |  | 199.3 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 0.7 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14

## 030
SHEATHED THERMOCOUPLE CALIBRATION RECORD | [17 empty] | 1/1 |  | Page
Management number |  |  |  | DVM - 172 - 030 | [8 empty] | Calibration day |  | 2017-11-02
Manufacturer - Model |  |  |  | DAIKI KOGYO - SKS 35 | [8 empty] | Expiration day |  | 2018-11-30
Equipment number |  |  |  | 103769 | [8 empty] | Effective time |  | 12 months since the last day of calibration month
Testing equipment
| Equipment name |  |  |  |  | Management number |  |  |  |  | Manufacturer - Model | [5 empty] | Equipment number
Standard thermocouple
| Standard digital thermometer |  |  |  |  | DVM - 008 - 001 |  |  |  |  | TATEYAMA KAGAKU - D619 | [5 empty] | 10100200149
1. Appearance test
| Corruption | [6 empty] | Result | [5 empty] | Evaluation criteria
|  | No |  | Yes |  |  |  |  | Pass |  |  | Fail |  | No damage
2. Insulation resistance test
| Insulation resistance value | [6 empty] | Result | [5 empty] | Evaluation criteria
| 1000 |  |  |  |  | MΩ |  |  | Pass |  |  | Fail |  | 100MΩ or more
3. Temperature test
| Test temperature |  |  | Indicated temperature of standard thermocouple (A) | [6 empty] | Indicated temperature of inspected thermocouple (B) |  |  |  |  | Difference (B-A)
| 100℃ |  |  | 99.2 |  |  |  |  | ℃ |  | 99.2 |  |  | ℃ |  | 0 |  |  | ℃
| 200℃ |  |  | 199.3 |  |  |  |  | ℃ |  | 200 |  |  | ℃ |  | 0.7 |  |  | ℃
| Result |  |  |  | Pass |  |  | Fail |  |  |  | Evaluation criteria |  | The difference is within ±2.5℃
[16 empty] | Checked by |  | Tested by
| Overall assessment |  |  |  |  | Pass
Fail
Retention period |  |  | 5 years |  |  |  | Storage department |  |  | Technical room |  |  | TEC-C-006E/V-002 |  |  | Revised (enacted) date |  | 2014-09-14