## Sheet1
Equipment | Type | Technical Requirement
pH Meter | Mobile type -0-14 | Range 0~14 pH / 0~±1999 mV, Accuracy : ±0,01 pH / ±1mV, Resistance Input : 1012J, Model HI8014 or equivalent
Electrical conductivity meter ( many measuring range ) | Mobile type - capable of measuring the electrical conductivity of pure and ultra- high range | Model HI8033 (HI 4321) or equivalent
Electronic balance scales placed above | Weight range from 0,01g to 3 kg | Max weight : 4100g, Accuracy : 0,01g, Model PAG4102 or equivalent
Magnetic stirrer | Volume 3L | Speed ​​100 ~ 1650 rev / min , 50R ~ 5y power , size 200 x 200 x 110 mm, Model HP-or equivalent
Desiccator | Volume 5L | Vacuum desiccator , 6L volumes , maximum load : 15kg , Size : 0286 x 354mm , Model VDR - 20 or equivalent
Glass thermometers | 0 - 100 C | Manufactured in Korea, according to the technical instructions or equivalent
Drying ovens up to 2000C | 50L | Capacity 60L, Room temperature: +10 ~ 250C, Control by PID with automatic rotation sensor with a K-CA, 2ea prices standard / maximum 8ea. Internal dimensions: 400x360x420 mm, Weight 53 kg. Model OF-02GW or equivalent
Glassware drying cabinet | Capacity 100L | 102L Volume, Amb +5 -700C, Controlled by P.I.D Auto Turning with PT 100 Type Sensor, Shelf Std. 2ea / Max. 12ea, Internal Size : 480x410x520 mm, Weight 61.5 kg, Model IB-15G or equivalent
Electric stove 2400 W | 400 x 200 mm | Up to 4000C, Heater 600W, Face 180x180 mm, Size 200x200x110 mm, or equivalent
Analytical balance | 1kg to 0,0001g | Capacity weight: 410g, Availability reading: 0.1 mg, Repeatability: 0.1 mg, Response Time: 2.5 seconds. Model ME414S or equivalent
Centrifugal machine | 0 to 6000 rpm | Maximum speed : 15,000 rpm, maximum RCF: 5,290 x g, maximum capacity:10 x 85 R, Size (CxRxD) 723x665x387mm, weight 45kg, Model Combi 514R or equivalent
Humidity Meter |  | Humidity Meter
Calculator | Science | CASIO, Model FX-4500PA or equivalent
Dissolved oxygen analyzer portable | Multi range - for the range mg / L and ug / L | Model HI9143 or equivalent
Stopwatch | Alarms | Manufactured in Korea , Quartz type , diameter of 350mm or equivalent
Buchholz gas tester |  | Manufactured in Korea, according to the technical instructions or equivalent
ORSAT gas analyzer |  | Model #39-540 or equivalent
Gas Analyzer Toolkit ( brand Draeger ) | For ammonia , chlorine gas , etc. | Model x-am7000 or equivalent
Hydrometer | SG ranging from 800 to 1500 , each of the hydrometer has 100 units ensures accuracy to 1 unit SG | Model DK - 701 or equivalent
Turbidity meter | Turbidity range from 0 to 1000 NTU | Model 2100Q or equivalent
Residual chlorine analyzer | Compare color machine , complete with tables DPD | Model DC-1200-CL or equivalent
Igniting temperature meter of the oil | -Determination of igniting temperature in the fuel , lubricants , fluids tend to produce membrane surface; | Conforms to ASTM D93 and related specifications; - It is possible to comply with the specifications of : AASHTO T73-811 , IP 34 , ISO 2719 , DIN 51 758 , FTM 791-1102 , NF F 07-019 ; - Power supply: 220V , 50Hz , single phase , 3.4a ; - Dimensions LxWxH (cm) : 24x20x57 ; - Weight : 10.9 kg
Portable gas analyzer | Equipment achieve the ISO 9001, ISO 14001 and OHSAS 18001 | Equipment achieve the ISO 9001, ISO 14001 and OHSAS 18001 - Simultaneous measurement of 3 parameters: NOx / SO2 / O2 - Measuring range: + NOx: 0 to 25/50/100/250/500/1000/2500 ppm + SO2: 0 to 200/500/1000 / 3000ppm + O2: 0 to 10/25% volume - Repeatability: ± 0.5% scale (NOx:> 100ppm) ± 1% on the scales if out of range - Linearity: ± 2% of scale - The drift: ± 1% of scale / day (with SO2: ± 2% of scale / day) - Response time (T90): 45 seconds or less, particularly SO2 is 180 seconds or less - Traffic Engineering samples: ~ 0.5 l / min - Screen: Touch, Display measured value (3 or 4 decimals), the measurement range, flow ... - Output: DC 4-20mA, Ethernet - Storage: SD memory card - Ambient temperature: 0 to 40oC - Humidity around 85% or less - Power supply: 220-240 V AC 50 / 60Hz - ~ Power consumption: 200VA - Size: 260W x 510D x 265H mm / no shield, 300W x 520D x 265H or mm / mm protective plate / protection plate - Weight (machine): 13 ~ 15kg - Conditions for sample gas: Temperature: <40 ° C + Water content: standard or less at ambient temperature. + Dust: 0.1 mg / m3 or less. + Pressure: ± 0,98kPa
Balance | Weight load 3kg; accuracy 0,01g | Model HF3000 or equivalent
Heat pump |  | Compliance with DIN 51 900 , ASTM 240D , ISO 1928 and BSI ; - Equipment easy to use, durable , are applications in the fields of raw materials extraction fuel ( coal , gas, oil ) ; - There are two methods of measurement is isoperibolic and dynamic pre-programmed and can be changed by the user . On the other hand , the working temperature can be selected for both methods depending on the conditions of the experiment .The quantitative decomposition can proceed in parallel with the analysis of total calories for the determination of halogens and sulfur; - Major equipment with screen and keyboard control activities and displays operating parameters; - The equipment is connected to the tank thermostat KV 600; - Devices with the ability to automatically fill oxygen in the sample cup decomposition; - You can upgrade a computer connected to the printer to process and print the results. 2.3. Key Specifications: - Working temperature 25 ° C or 30 ° C; - The range of the maximum analysis: 40.000J (equivalent to 9553 calories); - The device operates according to two methods: Isoperibolic and Dynamic in two temperature values ​​of 25 and 30 ° C; - Resolution: 0.0001 ° C; - The method of analysis Isoperibol time is 22 minutes; - According to the method Dynamic analysis time is 7 minutes (every hour measured about 8 samples); - Oxygen pressure is 30bar operation; - Flow rate of 60 liters / hour; - Relative humidity of 80%; - Power: 1800W; - Voltage: 220 ÷ 230V / 50Hz; - Dimensions (WxHxD): 440 x 450 x 500 mm; - Weight: 35kg
Lubricating oil viscometer |  | ' - The equipment is designed to determine the viscosity of the oil in accordance with ASTM D445 , ... can satisfy the standard D446 , IP17 , and more ... - Used to determine the viscosity of lubricating oil , FO , crude oil , and more ... - The ability to measure the 7 samples . - PID temperature control stability achieved quickly , avoid thermal shock . - Simultaneous display set temperature and the temperature tank.