## O2 - 10%
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | O2 | [8 empty] | Calibration range |  |  | R |  | 10 |  |  | %
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 0%
| Sensitivity adjustment gas | [6 empty] | CV19 |  |  |  |  | TAKACHIHO |  |  |  | 2BIS76220 |  |  |  | 9.151%
| Checking gas | [6 empty] | CV19 |  |  |  |  | TAKACHIHO |  |  |  | 2BIS76220 |  |  |  | 9.151%
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 9.18 | [9 empty] | 9.16 | [5 empty] | 100.2183 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
%
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 9.15 | [9 empty] | 9.14 | [5 empty] | -0.1 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | O2 | [8 empty] | Calibration range |  |  | R |  | 10 |  |  | %
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 0 |  |  |  |  | 0 |  |  | 0 |  |  | 9.14 |  |  | 0 |  |  | 0
| Second time |  |  |  | 0 |  |  |  |  | 0 |  |  | 0 |  |  | 9.14 |  |  | 0 |  |  | 0
| Third time |  |  |  | 0 |  |  |  |  | 0 |  |  | 0 |  |  | 9.14 |  |  | 0 |  |  | 0
| Average value |  |  |  | 0 |  |  |  |  | ̶ |  |  | 0 |  |  | 9.14 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## O2 - 25%
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | O2 | [8 empty] | Calibration range |  |  | R |  | 25 |  |  | %
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 0%
| Sensitivity adjustment gas | [6 empty] | CV19 |  |  |  |  | TAKACHIHO |  |  |  | 2BIS76220 |  |  |  | 9.151%
| Checking gas | [6 empty] | CV16 |  |  |  |  | TAKACHIHO |  |  |  | 4J21479 |  |  |  | 23.07%
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 18.7 | [9 empty] | 18.72 | [5 empty] | 99.8932 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
%
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 9.15 | [9 empty] | 9.14 | [5 empty] | -0.04 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | O2 | [8 empty] | Calibration range |  |  | R |  | 25 |  |  | %
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 0.03 |  |  |  |  | 0 |  |  | 0 |  |  | 23.07 |  |  | 0 |  |  | 0
| Second time |  |  |  | 0.02 |  |  |  |  | -0.01 |  |  | -0.04 |  |  | 23.07 |  |  | 0 |  |  | 0
| Third time |  |  |  | 0.03 |  |  |  |  | 0 |  |  | 0 |  |  | 23.08 |  |  | 0.01 |  |  | 0.04
| Average value |  |  |  | 0.03 |  |  |  |  | ̶ |  |  | -0.01 |  |  | 23.07 |  |  | ̶ |  |  | 0.01
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## SO2 - 200ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | SO2 | [8 empty] | Calibration range |  |  | R |  | 200 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 0.00ppm
| Sensitivity adjustment gas | [6 empty] | CV 23 |  |  |  |  | TAKACHIHO |  |  |  | 4J18183 |  |  |  | 90.02ppm
| Checking gas | [6 empty] | CV 23 |  |  |  |  | TAKACHIHO |  |  |  | 4J18183 |  |  |  | 90.02ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 90.1 | [9 empty] | 90.2 | [5 empty] | 99.8891 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 90.02 | [9 empty] | 89.92 | [5 empty] | -0.05 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | SO2 | [8 empty] | Calibration range |  |  | R |  | 200 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | -0.02 |  |  |  |  | 0.19 |  |  | 0.095 |  |  | 89.58 |  |  | 0.17 |  |  | 0.085
| Second time |  |  |  | -0.3 |  |  |  |  | -0.09 |  |  | -0.045 |  |  | 89.4 |  |  | -0.01 |  |  | -0.005
| Third time |  |  |  | -0.32 |  |  |  |  | -0.11 |  |  | -0.055 |  |  | 89.26 |  |  | -0.15 |  |  | -0.075
| Average value |  |  |  | -0.21 |  |  |  |  | ̶ |  |  | 0 |  |  | 89.41 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## SO2 - 500ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | SO2 | [8 empty] | Calibration range |  |  | R |  | 500 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 100%
| Sensitivity adjustment gas | [6 empty] | CV 23 |  |  |  |  | TAKACHIHO |  |  |  | 4J18183 |  |  |  | 90.02ppm
| Checking gas | [6 empty] | CV 23 |  |  |  |  | TAKACHIHO |  |  |  | 4J18183 |  |  |  | 90.02ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 89.45 | [9 empty] | 89.45 | [5 empty] | 100 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 90.02 | [9 empty] | 90 | [5 empty] | -0.004 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | SO2 | [8 empty] | Calibration range |  |  | R |  | 500 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | -0.65 |  |  |  |  | 0.12 |  |  | 0.024 |  |  | 89.45 |  |  | 0.03 |  |  | 0.006
| Second time |  |  |  | -0.8 |  |  |  |  | -0.03 |  |  | -0.006 |  |  | 89.45 |  |  | 0.03 |  |  | 0.006
| Third time |  |  |  | -0.85 |  |  |  |  | -0.08 |  |  | -0.016 |  |  | 89.35 |  |  | -0.07 |  |  | -0.014
| Average value |  |  |  | -0.77 |  |  |  |  | ̶ |  |  | 0 |  |  | 89.42 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## SO2 - 1000ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | SO2 | [8 empty] | Calibration range |  |  | R |  | 1000 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 0.0ppm
| Sensitivity adjustment gas | [6 empty] | CV22 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS31151 |  |  |  | 963.8ppm
| Checking gas | [6 empty] | CV22 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS31151 |  |  |  | 963.8ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 1000 | [9 empty] | 1001.3 | [5 empty] | 99.8702 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 963.8 | [9 empty] | 963 | [5 empty] | -0.08 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | SO2 | [8 empty] | Calibration range |  |  | R |  | 1000 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 0.4 |  |  |  |  | -0.13 |  |  | -0.013 |  |  | 961.4 |  |  | 1.4 |  |  | 0.14
| Second time |  |  |  | 1.2 |  |  |  |  | 0.67 |  |  | 0.067 |  |  | 960.1 |  |  | 0.1 |  |  | 0.01
| Third time |  |  |  | 0 |  |  |  |  | -0.53 |  |  | -0.053 |  |  | 958.5 |  |  | -1.5 |  |  | -0.15
| Average value |  |  |  | 0.53 |  |  |  |  | ̶ |  |  | 0 |  |  | 960 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## NO - 50ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NO | [8 empty] | Calibration range |  |  | R |  | 50 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 0.00ppm
| Sensitivity adjustment gas | [6 empty] | CV9 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS58094 |  |  |  | 45.17ppm
| Checking gas | [6 empty] | CV9 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS58094 |  |  |  | 45.17ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 44.02 | [9 empty] | 44.12 | [5 empty] | 99.7733 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 45.17 | [9 empty] | 45.2 | [5 empty] | 0.06 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NO | [8 empty] | Calibration range |  |  | R |  | 50 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | -0.02 |  |  |  |  | 0 |  |  | 0 |  |  | 45.19 |  |  | 0.05 |  |  | 0.1
| Second time |  |  |  | -0.03 |  |  |  |  | -0.01 |  |  | -0.02 |  |  | 45.13 |  |  | -0.01 |  |  | -0.02
| Third time |  |  |  | -0.02 |  |  |  |  | 0 |  |  | 0 |  |  | 45.09 |  |  | -0.05 |  |  | -0.1
| Average value |  |  |  | -0.02 |  |  |  |  | ̶ |  |  | -0.01 |  |  | 45.14 |  |  | ̶ |  |  | -0.01
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## NO - 100ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NO | [8 empty] | Calibration range |  |  | R |  | 100 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 0.00ppm
| Sensitivity adjustment gas | [6 empty] | CV8 |  |  |  |  | TAKACHIHO |  |  |  | 2BIS12853 |  |  |  | 89.73ppm
| Checking gas | [6 empty] | CV8 |  |  |  |  | TAKACHIHO |  |  |  | 2BIS12853 |  |  |  | 89.73ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 87.63 | [9 empty] | 88.01 | [5 empty] | 99.5682 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 89.73 | [9 empty] | 89.83 | [5 empty] | 0.1 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NO | [8 empty] | Calibration range |  |  | R |  | 100 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 0.05 |  |  |  |  | -0.01 |  |  | -0.01 |  |  | 89.45 |  |  | 0.02 |  |  | 0.02
| Second time |  |  |  | 0.05 |  |  |  |  | -0.01 |  |  | -0.01 |  |  | 89.42 |  |  | -0.01 |  |  | -0.01
| Third time |  |  |  | 0.07 |  |  |  |  | 0.01 |  |  | 0.01 |  |  | 89.42 |  |  | -0.01 |  |  | -0.01
| Average value |  |  |  | 0.06 |  |  |  |  | ̶ |  |  | 0 |  |  | 89.43 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## NO - 250ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NO | [8 empty] | Calibration range |  |  | R |  | 250 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 0.00ppm
| Sensitivity adjustment gas | [6 empty] | CV8 |  |  |  |  | TAKACHIHO |  |  |  | 2BIS12853 |  |  |  | 89.73ppm
| Checking gas | [6 empty] | CV6 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS58124 |  |  |  | 225.9ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 221.7 | [9 empty] | 222.3 | [5 empty] | 99.7301 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 89.73 | [9 empty] | 90.3 | [5 empty] | 0.228 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NO | [8 empty] | Calibration range |  |  | R |  | 250 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 0.35 |  |  |  |  | 0.04 |  |  | 0.016 |  |  | 225.45 |  |  | 0.07 |  |  | 0.028
| Second time |  |  |  | 0.22 |  |  |  |  | -0.09 |  |  | -0.036 |  |  | 225.3 |  |  | -0.08 |  |  | -0.032
| Third time |  |  |  | 0.35 |  |  |  |  | 0.04 |  |  | 0.016 |  |  | 225.38 |  |  | 0 |  |  | 0
| Average value |  |  |  | 0.31 |  |  |  |  | ̶ |  |  | 0 |  |  | 225.38 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## NO - 500ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NO | [8 empty] | Calibration range |  |  | R |  | 500 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 0.0ppm
| Sensitivity adjustment gas | [6 empty] | CV6 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS58124 |  |  |  | 225.9ppm
| Checking gas | [6 empty] | CV4 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS49168 |  |  |  | 442.1ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 430 | [9 empty] | 431.3 | [5 empty] | 99.6986 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 225.9 | [9 empty] | 226.2 | [5 empty] | 0.06 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NO | [8 empty] | Calibration range |  |  | R |  | 500 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 0.3 |  |  |  |  | -0.03 |  |  | -0.006 |  |  | 442 |  |  | -0.03 |  |  | -0.006
| Second time |  |  |  | 0.2 |  |  |  |  | -0.13 |  |  | -0.026 |  |  | 442.1 |  |  | 0.07 |  |  | 0.014
| Third time |  |  |  | 0.5 |  |  |  |  | 0.17 |  |  | 0.034 |  |  | 442 |  |  | -0.03 |  |  | -0.006
| Average value |  |  |  | 0.33 |  |  |  |  | ̶ |  |  | 0 |  |  | 442.03 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## NO - 2500ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NO | [8 empty] | Calibration range |  |  | R |  | 2500 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 0.0ppm
| Sensitivity adjustment gas | [6 empty] | CV2 |  |  |  |  | TAKACHIHO |  |  |  | 5J14765 |  |  |  | 1908ppm
| Checking gas | [6 empty] | CV2 |  |  |  |  | TAKACHIHO |  |  |  | 5J14765 |  |  |  | 1908ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 2022 | [9 empty] | 2028 | [5 empty] | 99.7041 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 1908 | [9 empty] | 1909 | [5 empty] | 0.04 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NO | [8 empty] | Calibration range |  |  | R |  | 2500 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 3 |  |  |  |  | 0.33 |  |  | 0.0132 |  |  | 1903 |  |  | 0.67 |  |  | 0.0268
| Second time |  |  |  | 2 |  |  |  |  | -0.67 |  |  | -0.0268 |  |  | 1900 |  |  | -2.33 |  |  | -0.0932
| Third time |  |  |  | 3 |  |  |  |  | 0.33 |  |  | 0.0132 |  |  | 1904 |  |  | 1.67 |  |  | 0.0668
| Average value |  |  |  | 2.67 |  |  |  |  | ̶ |  |  | 0 |  |  | 1902.33 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## NOx - 50ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NOx | [8 empty] | Calibration range |  |  | R |  | 50 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 0.00ppm
| Sensitivity adjustment gas | [6 empty] | CV9 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS58094 |  |  |  | 45.17ppm
| Checking gas | [6 empty] | CV9 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS58094 |  |  |  | 45.17ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 44.23 | [9 empty] | 44.51 | [5 empty] | 99.3709 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 45.17 | [9 empty] | 45.28 | [5 empty] | 0.22 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NOx | [8 empty] | Calibration range |  |  | R |  | 50 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 0.03 |  |  |  |  | 0 |  |  | 0 |  |  | 45.08 |  |  | -0.04 |  |  | -0.08
| Second time |  |  |  | 0.03 |  |  |  |  | 0 |  |  | 0 |  |  | 45.15 |  |  | 0.03 |  |  | 0.06
| Third time |  |  |  | 0.02 |  |  |  |  | -0.01 |  |  | -0.02 |  |  | 45.12 |  |  | 0 |  |  | 0
| Average value |  |  |  | 0.03 |  |  |  |  | ̶ |  |  | -0.01 |  |  | 45.12 |  |  | ̶ |  |  | -0.01
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## NOx - 100ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NOx | [8 empty] | Calibration range |  |  | R |  | 100 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 100%
| Sensitivity adjustment gas | [6 empty] | CV8 |  |  |  |  | TAKACHIHO |  |  |  | 2BIS12853 |  |  |  | 89.73ppm
| Checking gas | [6 empty] | CV8 |  |  |  |  | TAKACHIHO |  |  |  | 2BIS12853 |  |  |  | 89.73ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 86.96 | [9 empty] | 87.21 | [5 empty] | 99.7133 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 89.73 | [9 empty] | 90.02 | [5 empty] | 0.29 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NOx | [8 empty] | Calibration range |  |  | R |  | 100 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 0.07 |  |  |  |  | 0 |  |  | 0 |  |  | 89.77 |  |  | -0.07 |  |  | -0.07
| Second time |  |  |  | 0.07 |  |  |  |  | 0 |  |  | 0 |  |  | 89.83 |  |  | -0.01 |  |  | -0.01
| Third time |  |  |  | 0.07 |  |  |  |  | 0 |  |  | 0 |  |  | 89.93 |  |  | 0.09 |  |  | 0.09
| Average value |  |  |  | 0.07 |  |  |  |  | ̶ |  |  | 0 |  |  | 89.84 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## NOx - 250ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NOx | [8 empty] | Calibration range |  |  | R |  | 250 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 100%
| Sensitivity adjustment gas | [6 empty] | CV8 |  |  |  |  | TAKACHIHO |  |  |  | 2BIS12853 |  |  |  | 89.73ppm
| Checking gas | [6 empty] | CV6 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS58124 |  |  |  | 225.9ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 223.4 | [9 empty] | 224.2 | [5 empty] | 99.6432 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 89.73 | [9 empty] | 89.72 | [5 empty] | -0.004 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NOx | [8 empty] | Calibration range |  |  | R |  | 250 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 0.1 |  |  |  |  | -0.02 |  |  | -0.008 |  |  | 225.23 |  |  | -0.12 |  |  | -0.048
| Second time |  |  |  | 0.1 |  |  |  |  | -0.02 |  |  | -0.008 |  |  | 225.3 |  |  | -0.05 |  |  | -0.02
| Third time |  |  |  | 0.17 |  |  |  |  | 0.05 |  |  | 0.02 |  |  | 225.52 |  |  | 0.17 |  |  | 0.068
| Average value |  |  |  | 0.12 |  |  |  |  | ̶ |  |  | 0 |  |  | 225.35 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## NOx - 500ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NOx | [8 empty] | Calibration range |  |  | R |  | 500 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 100%
| Sensitivity adjustment gas | [6 empty] | CV6 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS58124 |  |  |  | 225.9ppm
| Checking gas | [6 empty] | CV4 |  |  |  |  | TAKACHIHO |  |  |  | 3BIS49168 |  |  |  | 442.1ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 425.3 | [9 empty] | 428.7 | [5 empty] | 99.2069 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 225.9 | [9 empty] | 227.1 | [5 empty] | 0.24 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NOx | [8 empty] | Calibration range |  |  | R |  | 500 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 0.3 |  |  |  |  | -0.13 |  |  | -0.026 |  |  | 442.2 |  |  | -0.7 |  |  | -0.14
| Second time |  |  |  | 0.5 |  |  |  |  | 0.07 |  |  | 0.014 |  |  | 443.5 |  |  | 0.6 |  |  | 0.12
| Third time |  |  |  | 0.5 |  |  |  |  | 0.07 |  |  | 0.014 |  |  | 443 |  |  | 0.1 |  |  | 0.02
| Average value |  |  |  | 0.43 |  |  |  |  | ̶ |  |  | 0 |  |  | 442.9 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01

## NOx - 2500ppm
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 1/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Manufacturer - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NOx | [8 empty] | Calibration range |  |  | R |  | 2500 |  |  | ppm
Checking Equipment
| Equipment name | [6 empty] | Management number |  |  |  |  | Manufacturer |  |  |  | Cylinder number |  |  |  | Concentration
| Zero adjustment gas | [6 empty] | CV26 |  |  |  |  | TAKACHIHO |  |  |  | 7J26824 |  |  |  | 100%
| Sensitivity adjustment gas | [6 empty] | CV2 |  |  |  |  | TAKACHIHO |  |  |  | 5J14765 |  |  |  | 1908ppm
| Checking gas | [6 empty] | CV2 |  |  |  |  | TAKACHIHO |  |  |  | 5J14765 |  |  |  | 1908ppm
| A/D converter |  |  |  | Management number | [5 empty] | DVM - 149 - 001 |  |  |  |  | Equipment number |  |  |  | 93310021
[5 empty] | Manufacturer - Model | [5 empty] | KEYENCE NR 500 |  |  |  |  | Expiration day |  |  |  | 2017-09-30
| Computer using | [9 empty] | 15170
1. Responsiveness
| Type of the target concentration meter | [9 empty] | NOx, O2, SO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 4 minutes | [9 empty] | After 10 minutes
| A | [9 empty] | B | [5 empty] | e
| 2038 | [9 empty] | 2053 | [5 empty] | 99.2694 | [5 empty] | %
| Type of the target concentration meter | [9 empty] | CO, CO2
| Indicated value of the span calibration gas | [15 empty] | Indication rate
| After 2 minutes 30 seconds | [9 empty] | After 6 minutes
| A | [9 empty] | B | [5 empty] | e
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Indication rate is more than 90%
Calculation formula
2. Instrument error test
| Testing gas | [15 empty] | Instrument error compares with the highest concentration
| Concentration | [9 empty] | Indicated value
| C | [9 empty] | I | [5 empty] | e
| 1908 | [9 empty] | 1912 | [5 empty] | 0.16 | [5 empty] | %
| Result | [5 empty] | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Instrumental error is within ± 4%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01
NOx/O2 ANALYZER CALIBRATION RECORD | [22 empty] | 2/2 | Page
Management number | [5 empty] | DVM - 039 - 003 | [8 empty] | Calibration day |  |  | 2017-05-31
Nhà sản xuất - Model | [5 empty] | HORIBA PG - 337 | [8 empty] | Expiration day |  |  | 2018-05-31
Equipment number | [5 empty] | KD65DEUG | [8 empty] | Effective time |  |  | 12 months since the last day of calibration month
Type of the target concentration meter | [5 empty] | NOx | [8 empty] | Calibration range |  |  | R |  | 2500 |  |  | ppm
3. Repeatability test
[5 empty] | ZERO adjustment gas | [10 empty] | Span calibration gas
[5 empty] | Indicated value |  |  |  |  | Average error |  |  | Error compares with the highest concentration |  |  | Indicated value |  |  | Average error |  |  | Error compares with the highest concentration
[5 empty] | C0 |  |  |  |  | δ0 |  |  | e0 |  |  | Cs |  |  | δs |  |  | es
|  |  | Unit | [9 empty] | % | [8 empty] | %
| First time |  |  |  | 3 |  |  |  |  | -0.33 |  |  | -0.0132 |  |  | 1914 |  |  | -4.67 |  |  | -0.1868
| Second time |  |  |  | 3 |  |  |  |  | -0.33 |  |  | -0.0132 |  |  | 1920 |  |  | 1.33 |  |  | 0.0532
| Third time |  |  |  | 4 |  |  |  |  | 0.67 |  |  | 0.0268 |  |  | 1922 |  |  | 3.33 |  |  | 0.1332
| Average value |  |  |  | 3.33 |  |  |  |  | ̶ |  |  | 0 |  |  | 1918.67 |  |  | ̶ |  |  | 0
| Result |  |  |  |  | Pass |  |  |  | Fail |  |  | Evaluation criteria |  |  | Error is within ± 2%
Calculation formula
[19 empty] | Checked by |  |  |  | Tested by
| Overall assessment | [5 empty] | Pass
Fail
Retention period |  |  |  | 5 years |  |  |  |  | Storage department |  |  | Measurement office | [6 empty] | Revised (Enacted) date |  |  |  | 2014-09-01