## mCERTs
Performance Characteristics for CWMs
| Temp | pH | Conduvtivity | Turbidity | DO | COD | TOC | Ammonia as N | NO3, NO2, TON as N | TP, o-phosphate
| ℃ | pH units | % of Span | % of Span | % of reading | % of reading | % of reading | % of reading | % of reading | % of reading
Mean Error | 0.3 | 0.2 | 1 | 2 (or 0.2 FTU) | 5 (or 0.2 mg/L or 2 % sat) | 10 (or 10 mg/L) | 10 (or 0.2 mg/L) | 10 (or 0.03 mg/L- as N) | 5 (or 0.2 mg/L- as N) | 10 (or 0.05 mg/L- as P)
Linearity | 0.2 | 0.1 | 0.2 | 1 | 2.5 | 5 | 5 | 5 | 2.5 | 5
Repeatibility | 0.2 | 0.1 | 0.5 | 1 (or 0.1 FTU) | 2.5 (or 0.1 mg/L or 1 % sat) | 5 (or 5 mg/L) | 5 (or 0.1 mg/L) | 5 (or 0.015 mg/L- as N) | 2.5 (or 0.1 mg/L- as N) | 5 (or 0.025 mg/L- as P)
Sample matrix effect | 0 | 0 | 0 | 0 | 2.5 | 5 | 3 % span | 5 | 2.5 | 5
Drift | 0.2 | 0.1 | 0.5 | 1 | 2.5 | 5 | 5 | 5 | 2.5 | 5
Voltage impedance | 0.2 | 0.05 | 0.5 | 0.5 | 1 | 2.5 | 2.5 | 2.5 | 1 | 2.5
Supply voltage | 0.2 | 0.05 | 0.25 | 0.5 | 1 | 2.5 | 2.5 | 2.5 | 1 | 2.5
Ambient temp., | 0.2 | 0.1 | 0.5 | 1 | 2.5 | 5 | 5 | 5 | 2.5 | 5
Relative humidity | 0.2 | 0.1 | 0.5 | 1 | 2.5 | 5 | 5 | 5 | 2.5 | 5
Incident light | 0.2 | 0.05 | 0.25 | 1 | 1 | 2.5 | 2.5 | 2.5 | 1 | 2.5
Sample temp., | 0 | 0.1 | 0.5 | 1 | 2.5 | 5 | 5 | 5 | 2.5 | 5
Sample flow rate | 0 | 0.05 | 0.2 | 0.5 | 1 | 2.5 | 2.5 | 2.5 | 1 | 2.5
Sample pressure | 0 | 0.05 | 0.2 | 0.5 | 1 | 2.5 | 2.5 | 2.5 | 1 | 2.5
Combined performance characteristic | 0.5 | 0.3 | 1.5 | 2.5 | 6 | 12 | 12 | 12 | 6 | 12
Response time
Initial warm up
Mean Error | For any individual test point, the error, x, is the difference between the value given by the CWM and the conventional true value
Repeatibility | Repeatability is standard deviation of the measurements taken at reference conditions
Drift | expressed as the maximum change in error over the period of the test.
* CWMs: continuous water monitors
Note on F8: A dissolved oxygen CWM shall be tested to check for the effect of chlorides present in the sample.
Note on A19: Where an instrument operates in a batch mode based on the development of a chemical reaction to the sample over a fixed period of time, the response time and warm-up drift tests may be omitted.
Note on A20: Where an instrument operates in a batch mode based on the development of a chemical reaction to the sample over a fixed period of time, the response time and warm-up drift tests may be omitted.

## KOREA
| pH | SS | DO | COD | BOD | TN | TP | TOC | Turbidity
| pH units | % of span
Measuring Range | 0 ~ 14 | 0 ~ 1,000 mg/L | 0 ~ 20 mg/L | 0 ~ 250 mg/L | 0 ~ 200 mg/L | 0 ~ 100 mg/L | 0 ~ 20 mg/L | 0 ~ 100 mg/L | 0 ~ 10 NTU (drinking water)
Measurement Principle | Glass or Antimony electrode | gravimetric analysis, light scattered method | Polarographic or galvanic (membrane) | KMnO4 | DO sensor | UV absorption or Cd-reduction | ISE or Absorption Spectroscopy (Ascorbic acid reduction) | combustion or wet chemical oxidation | light scattered method (Nephelometry)
Detection Limit | 0.1 | 1 mg/L | 0.1 mg/L | 1 mg/L | 1 mg/L | 1 mg/L | 0.05 mg/L | 0.1 mg/L | 0.1 NTU (drinking water)
Repeatability | 0.1 | 3% | 0.3 mg/L | 3% | 3% | 3% | 3% | 3% | 2%
Span Drift | 0.1 | 5% | 0.2 mg/L | 5% | 5% | 5% | 5% | 5% | 3%
Zero Drift | 0.1 | 5% | 0.3 mg/L | 5% | 5% | 5% | 5% | 5% | 3%
Linearity | ± 0.1 | ± 5% | - | ± 5% | - | ± 5% | ± 5% | ± 5% | ± 5%
Response Time | 30 sec (t90) | - | 2 min(t90) | - | - |  |  | 15 min(t90) | 10 min(t90)
Glucose test | COD: ± 5 %; BOD: -
Interference test | DO: ± 0.3 mg/L; TOC: 5%
Recovery test | TOC: 90%
Temp., compensation | ± 0.1 | - | ± 0.3 mg/L | - | -
Equivalence Input | pH: ± 0.1; SS: -; DO: -; COD: -
Voltage Regulation | 0.1 |  | 0.2 mg/L | 3% | 3% | 3% | 3% | 3%
Insulation resistance | > 2 ㏁ | > 2 ㏁ | > 2 ㏁ | > 2 ㏁ | > 2 ㏁ | > 2 ㏁ | > 2 ㏁ | > 2 ㏁ | > 2 ㏁
Withstanding voltage test | should be normal | should be normal | should be normal | should be normal | should be normal | should be normal | should be normal | should be normal | should be normal
RelativeAccuracy | ±10 % | ±30 % | ±20 % | ±20 % | ±30 % | ±20 % | ±20 % | ±10 %
Performance test (168 h) | Zero driff: 0.2 pH | Zero driff: 7 % | Zero driff: 0.3 mg/L | Zero driff: 7 % | Zero driff: 7 % | Zero driff: 7 % | Zero driff: 7 % | Zero driff: 7 % | Zero driff: 5 %
| Span drift: 0.2 pH | Span drift: 7 % | Span drift: 0.5 mg/L | Span drift: 7 % | Span drift: 7 % | Span drift: 7 % | Span drift: 7 % | Span drift: 7 % | Span drift: 5 %

## Index ...EPA Method
Method Number Index to EPA Test Methods
Parameter | Method | EPA Number
Temperature | Thermometric | 170.1
pH | Electrometric | 150.1
| Continuous Monitoring (Electrometric) | 150.2
SS | Gravimetric | 160.1 ~ 160.4
DO | Membrane Electrode (Probe) | 360.1
COD | Colorimetric | 410.4
| Semi-Automated Colorimetric | 410.4
| Titrimetric, High Level | 410.3
| Titrimetric, Low Level | 410.2
| Titrimetric, Mid- Level | 410.1
TOC | Combustion or Oxidation | 415.1
| UV Promoted | 415.2
Ammonia as N | Colorimetric | 350.1
| Semi-Automated Colorimetry | 350.1
| Colorimetric, Titrimetric | 350.2
| Potentiometric | 350.3
Nitrate (as nitrogen) | Colorimetric, Brucine | 352.1
Nitrogen | Nitrate-Nitrite - Colorimetric/Hydra | 353.1
Nitrate-Nitrate | by Automated Colorimetry | 353.2
Nitrate-Nitrite (as N) | Colorimetric/Cadmium | 353.2
| Manual Cadmium | 353.3
Nitrate (as Nitrogen) | Spectrophotometric | 363.6
orthphosphate (as P) | Automated Colorimetry | 365.1
Phosphorus | All Forms - Colorimetric/Automated | 365.1
| manual colorimetric/1 reagent | 365.2
| colorimetric/ 2 reagents | 365.3
| Total - Colorimetric/Automated | 365.4
| Orthophosphate by automated colorimetric | 365.5
Turbidity | Nephelometric | 180.1