CTA solution

# Facility

- Please describe the facility and include a sketch of the test section dimensions

Answer:

Dust particle mixing tower: height (2090 ± 0.003) mm, inner diameter of pipe is (160 ± 0.002) mm, in the middle is separated by insulating material

- Where do you want to measure in the facility

  - Area or volume to be measured (dimensions)?

Answer:

Area or volume to be measured for the anemometer as above

  - Do you need a traversing mechanism? (if there is one already, what type of traversing mechanism do you have: manual, software controlled?)

Answer:

Yes, Optimal with controller

  - How is the physical access into the test section (slit, hole etc)? What is the size, shape, and position of the access window?

Answer:

Test section by hote. The minimum size of the window is to be able to put an entire device inside (20 ~ 25 cm x 20 ~ 25 cm)

- Please describe environmental conditions of the facility

  - Temperature

  - Humidity

  - Pressure

  - Dust

  - Explosion hazard

  - Other, please specify

Answer:

Envirommental condition:

  - Temperature: (23  5) oC

  - Humidity: (40 ÷ 80) %RH.

  - Pessure: (600 ÷ 1150) hPa

  - Dust: ISO 12301-1 A1 and A2

  - Other, please specify: No chemical corrosive substances; no vibration and interference

# Experimental goal

- What is to be measured? Velocity, Temperature, Humidity etc...

Answer:

Measured: Velocity and controller condition (temperature; humidity; pressure)

- Will the analysis performed by the system and results presented by the software?

Answer:

Yes, by the system and results presented by the software or Data is extracted to a file for processing and analysis

- Or exchange data with other application (e.g. MATLAB, TecPlot), please specify

Answer:

MATLAB; excel

- Please provide publications on similar experiments in similar facilities, if any

Answer:

Facility for calibration of optical and condensation particle counters based on a turbulent aerosol mixing tube and a reference optical particle counter.

https://doi.org/10.1063/1.5095853

- Please specify the medium

  - Water

  - Air

  - Other, please specify

  - How clean is the medium?

Answer: Air

- Which velocity component(s)? Please indicate definition/co-ordinate system on the sketch

Answer:

Only velocity, no-definition/co-ordinate system

- If flow contains particles, what is the material state (solid or liquid) and what is the size range?

Answer:

Material: solid (size 2 -10 µm)

- What is the required distance from the sensor to the electronics?

Answer:

Distance maximum 30 cm

- Can the user be close to the equipment or must the system be controlled from a separate room? In the latter case, what is the distance from the setup to the control room?

Answer:

The user be close to the equipment

- Do you need a velocity / directional calibration device for the hotwire probes?

  - Manual

  - Automatic

Answer:

Yes, manual

- Are measurements related to a cyclic phenomenon

  - Once per revolution/period pulse available

  - One pulse per ____ degrees rotation or per  ____ msec available

Answer:

No

- Must control experiment from other system, please specify

Answer:

When the measured data from the system (display) is not correct with the set value, it will have to adjust the inlet flow through the pump so that the displayed speed approximates the set speed.

- Is the system for educational use or for research purposes?

Answer:

System use for reseach purposes

# Measurement requirements

- Flow mapping in a point, an area or a volume?

  - Required spatial resolution?

Answer:

Flow mapping a volume with diameter 160 mm; velocities between form 0.15 m/s to 0.5 m/s

- Dynamical analysis (turbulence spectra, time scales) in one or a few positions

  - Required temporal resolution?

Answer:

It is necessary to determine the uniformity of gas flow velocity in a 160 mm diameter pipe with time

- Please provide the properties of the flow

  - Velocity range

  - Temperature range

  - Humidity range

  - Pressure range

  - Other relevant information, please specify

Answer:

  - Velocity range: (0,15 ÷ 50) m/s; accuracy: ≤ 2%

  - Temperature range: (23  5) oC

  - Humidity range: (40 ÷ 80) %RH

  - Pressure range: (600 ÷ 1150) hPa

  - Other relevant information, please specify

# Process

- Will there be a tender process?

Answer:

Yes

- Is training required?

Answer:

Yes, receive training and instructions on how to use the equipment

- What is the required delivery time?

Answer:

before June, 2022

- What is the available budget?

Answer:

1000 – 1200 USD

Sales Qualification Guidance 0537_v1. Subject to change without notice. Copyright © 2017. Dantec Dynamics.

All Rights Reserved. www.dantecdynamics.com