## Table 1
Ecotech Document Control
EMS 0371-S40
S40 NOx Analyser Calibration Report
Notes:
Calibration Performed by |  |  | Khanh Nguyen
Date |  |  | 2028-08-13
Time Begin/End |  |  | 12:00 |  | 18:00
Location |  |  | Vinh Yen
Date:
2013-03-04
Customer | Thanh Thien
Instrument | Serinus 40
ID No. | 13-0472
System/Job No. | 500926
Calibration Equipment
Calibrator Model |  |  | Gascal 1100
ID/Serial No. |  |  | 13-0464
Zero Air Source Model |  |  | 8310LC
ID/Serial No. |  |  | 13-0466
Gas Std. Cylinder Serial No. |  |  |  | GTA060090
Cylinder Concentration (ppm) |  |  |  | 59.2
Cylinder Expiry Date |  |  |  | 2019-07-14
Cylinder Pressure (Bar) |  |  |  | 120
Cylinder Calibration |  |  |  | PASS
Displayed Instrument Parameters
Temperatures Pressures and Flow
Voltages Pots
Cell Temp |  | 49.9 - 50.1 |  |  | 50.04
Converter Temp |  | 320 - 330 |  |  | 325.3
Chassis Temp |  | 15 - 45 |  |  | 36.9
PMT |  | 12.95 - 13.05 |  |  | 13
Manifold Temp |  | 49.9 - 50.1 |  |  | 49.99
Ambient |  |  |  | 752.16
Cell | 70 - 200 |  |  | 164.1
Manifold |  |  |  | 705
Manifold 2 |  |  |  | 714.34
Flow | 0.6 - 0.66 |  |  | 0.9
High Voltage |  | 640 - 670 |  |  | 659.5
Analogue Supply |  | 11.5 - 12.5 |  |  | 12
Digital Supply |  | 4.85 - 5.15 |  |  | 5
-10 V |  | -9.5 - 10.5 |  |  | -9.9
HV Adjust | 130 - 150 |  |  | 141
PGA Gain | 1 - 128 |  |  | 16
Input Pot | 104 |  |  | 104
Test Pot | 0 |  |  | 0
Instrument Units |  |  | ppb
Firmware Version 2.09.0005
Single Point Calibration
Initial Span Instrument Gain |  |  |  | 0.87
Initial NO Offset (ppb) |  |  |  | 0
Initial NO2 Offset (ppb) |  |  |  | 0
Full Scale (ppb) 500
Precalibration Check
Calibration Point | Zero flow (sccm) |  |  | Span flow (sccm) |  | Calculated NOx (ppb) | Measured (ppb) |  |  |  | Error (%FS)
[7 empty] | NO | NO2 |  | NOx
Zero | 3970 |  |  | 0 |  | 0 | 0.37 | 0.73 |  | 1.1 | 0.2%
80% FS | 4441 |  |  | 30.38 |  | 402 | 485.64 | -1.55 |  | 484.2 | 16.4%
Note: Pre calibration check shall be within 5% of FS value for NOx span
Zero values should be with in 2% of Full Scale
Complete Filter change and perform leak check
Post-calibration Check
PASS/ FAIL FAIL
80% FS | 4441 |  |  | 30.38 |  | 402 | 397.87 | 2.26 |  | 400.19 | -0.4%
Zero | 3970 |  |  | 0 |  | 0 | 0.32 | 0.65 |  | 0.97 | 0.2%
Note: Do not fill in shaded cells Note: Post calibration check shall be within 1% of FS value for both NOx span and all zeros
PASS/ FAIL PASS
Final Span Instrument Gain |  |  |  | 0.96
Final NO Offset (ppb) |  |  |  | 0
Final NO2 Offset (ppb) |  |  |  | 0
Page 1
Multipoint Linearity Check
Calibration Point | Zero flow (sccm) |  |  | Span flow (sccm) |  | Calculated NOx (ppb) | Expected NOx (ppb) |  | Measured NOx (ppb) |  | Error (%)
Zero | 3970 |  |  | 0 |  | 0 | 0 |  | 0.97
20% FS | 4464 |  |  | 7.58 |  | 100 | 100 |  | 98.82 |  | -1.2%
40% FS | 4457 |  |  | 15.18 |  | 201 | 201 |  | 202.01 |  | 0.5%
60% FS | 4451 |  |  | 22.78 |  | 301 | 301 |  | 304.3 |  | 1.1%
80% FS | 4441 |  |  | 30.38 |  | 402 | 402 |  | 400.19 |  | -0.5%
Zero | 3970 |  |  | 0 |  | 0 | 0 |  | 0.53
Note: Do not fill in the shaded cells
450
400
350
300
250
y = 0.9989x + 0.675
R² = 0.9998
200
150
100
50
0
0 50 100 150 200 250 300 350 400 450
Expected Concentration
Standard error about y = x (% FS) |  |  |  | 0.47%
R-Square Value |  |  |  | 0.9998
Intercept |  |  |  | 0.675
PASS/ FAIL PASS
Note: Final Zero must be within 1% of FS,
Standard error of y about x must be < 1%FS
Technicians Signature Date
Page 2