Guidelines for creating calibration forms

- All font should be Calibri and its size should be 11, except form title and note

- Note font size should be 9

Title color

- Date format should be

- Time format should be

- All headers or banners should be bald and filled with color

Header or banner color

- All forms should include pre-calibration and  post calibration

- Try and use the same formula setup and layout for the cylinder calibration pass/fail section (based on date) on all forms

- All PASS/FAIL sections should be locked. Use conditional formatting to make the colors change for Pass/Fail statement

- For “PASS” condition color

- For “FAIL” condition color

- For “N/A” condition color

- All formula cells should be locked and filled with color.

Formula cell color

- Try and make formula cells go blank when not used rather than say Division by Zero or other error messages. N/A is fine of course.

- Use the below formulas for error and concentration calculation

- Calculated concentration = (Span Flow x Gas bottle concentration (ppb)) ÷ (Span Flow + Zero flow)

- Error is: (measured - calculated)

- Error as % Point is: (measured-calculated) ÷ calculated x 100

- Error as % of full scale is: (measured-calculated) ÷ (instrument full scale) x 100

- Use the below formula for standard error, R-square and intercept calculation

Standard error = (STEYX(Y_values,X_values))/full scale value

R-square value = RSQ(Y_values, X_values)

Intercept value = INTERCEPT(Y_values,X_values)

Please note STEYX, RSQ and INTERCEPT functions are built in EXCEL

- For single point, ensure the Pass/fail criteria for each gas is matched our standard gas analysers

Eg  Nox

- Pre calibration check shall be within 5% of FS value for NOx span

- Zero values should be within 2% of Full Scale - Post calibration check shall be within 1% of FS value for both NOx span and all zeros

- For the multi-point, ensure the PASS/FAIL criteria for each gas is matched our standard gas analysers

Eg Nox

Final Zero must be within 1% of FS,

Standard error of y about x must be < 1%FS

- Ensure the number of points required for each gas for multipoint is matched our standard

- The page layout should have the same margins as other forms.  The page percentage scaling should be 100% wherever possible.

- All text label alignment should be set to Left or center and bottom

- Use spin control for listing

- Signature and date section should be like this.

- Keep all columns to the same size as possible (do not manually resize column). Use  merge feature to create bigger column area

- Make sure you check how it looks when you print it off.  Ensure all modified /edited cells are within printed area.  Use print preview to check

- Try and make the form flows so that technicians do things in the correct order

- Try keeping the form as neat and clear as possible.  Add things that you believe they are important

Caroline